Procedure:  Alethia™ Group A Streptococcus

Procedure and revision number: SN11022 CLSI 05/20

Meridian Catalog #480150

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Alethia Group A Streptococcus (Group A Strep) assay, performed on the Alethia reader, is a qualitative in vitro diagnostic test for the detection of Streptococcus pyogenes (Group A β-hemolytic Streptococcus) in throat swab specimens.

The Alethia Group A Strep assay utilizes loop-mediated isothermal DNA amplification (LAMP) technology to detect Streptococcus pyogenes by targeting a segment of the Streptococcus pyogenes genome.  Results from the Alethia Group A Strep assay can be used as an aid in the diagnosis of Group A Streptococcal pharyngitis.  The assay is not intended to monitor treatment for Group A Streptococcus infections.

Alethia Group A Strep is intended for use in hospital, reference or state laboratory settings.

Group A Streptococcus, or Streptococcus pyogenes, is a bacterium commonly found in the human throat or on skin.  S. pyogenes causes a wide variety of diseases in humans, the most common being acute pharyngitis or strep throat.  Pharyngitis is diagnosed in approximately 11 million patients in the United States each year, with Group A Streptococcus accounting for 15% - 30% of cases in children and 5% - 20% of cases in adults.  Streptococcal pharyngitis is often accompanied by sore throat, fever, tonsillar exudates  and enlarged cervical lymph nodes.

Group A Streptococcus infections may result in mild illness (e.g. pharyngitis, impetigo) or may lead to invasive, life-threatening illness (e.g. cellulitis, bacteremia, necrotizing fasciitis, Streptococcal toxic shock syndrome).  Approximately 25% of patients with necrotizing fasciitis will die from the infection; more than 35% of patients with streptococcal toxic shock syndrome will die from the infection.  Streptococcus pyogenes may also present in healthy, asymptomatic patients.  Throat culture surveys of healthy children taken during school outbreaks have shown a carrier prevalence of up to 20%.

Signs and symptoms of Group A Streptococcus pharyngitis often overlap with other infections, making laboratory diagnosis an important tool in identification and treatment of infection.  Specimens for laboratory testing should be obtained by vigorous swabbing of both tonsils and the posterior pharynx.  Traditional laboratory diagnosis is performed by rapid antigen testing or culture on sheep blood agar followed by differentiation with latex agglutination.  Specificities of rapid antigen tests is variable and confirmation of negative results with follow-up throat swab culture is recommended.  Throat swab cultures may be 90 to 95% sensitive when correct sampling and plating techniques are used.

Diagnostic tests such chemilumiscent DNA probes may be as sensitive as standard throat cultures on sheep blood agar.

SPECIMEN:

Preferred Sample Type:

- Throat swabs in Liquid Amies or Stuart’s transport media.

- Acceptable swabs:  Rayon, polyester, or flocked nylon with plastic shafts.

Undesirable Specimens:

- Do not use throat swabs with Liquid Amies with charcoal transport media.

- Eswabs

- Collection and Preparation: Throat sample collection should be performed in accordance with institutional guidelines for collection of clinical specimens for culture of Group A Streptococcus.  Samples should be collected by vigorously swabbing the tonsils and the posterior pharynx.

- Place swab(s) in a non-nutritive transport medium (e.g. Liquid Amies, without charcoal or Liquid Stuart) and transport to the laboratory.  Samples should be held at 2-27 C during transport.

Samples may be held at 21-27 C for up to 48 hours prior to testing.  When testing will not be initiated within this time, the sample may be stored at 2-8 C for up to seven days.

This facility’s procedure for specimen collection is: _____________________________________ ______________________________________________________________________________

This facility’s procedure for transporting specimens is: __________________________________ ______________________________________________________________________________

This facility’s procedure for rejected specimens is: _____________________________________ ______________________________________________________________________________

MATERIALS AND EQUIPMENT:

- REAGENTS/MATERIALS PROVIDED

- The maximum number of tests obtained from this test kit is listed on the outer box.

- Alethia Sample Preparation Apparatus II/Negative Control III (SMP PREP): Tris-buffered solution containing formalin-treated E. coli harboring S. aureus DNA and sodium azide (0.09%) as a preservative.

- Alethia Group A Strep Test Device: Two-chambered device containing lyophilized amplification reagents (DNA polymerase, deoxynucleotide triphosphates) and either Group A Streptococcus-specific primers (TEST Chamber) or Staphylococcus aureus-specific primers (CONTROL Chamber).

- Alethia Heat Treatment Tubes

MATERIALS PROVIDED SEPARATELY

Alethia Group A Streptococcus (Group A Strep) External Control Kit, Catalog Number: 479910

- MATERIALS NOT PROVIDED

- Disposable latex gloves, powder free

- DNase/RNase-free, aerosol resistant pipette tips

- Specimen collection and transport system

- Swabs, breakable plastic shaft: Rayon, Polyester or Flocked Nylon

- Non-nutritive Transport Medium: Liquid Amies, without charcoal; Liquid Stuart

EQUIPMENT NOT PROVIDED:

- Dry-bath with 12 mm heat block capable of 95 C

- Digital thermometer with Max/Min Temperature Memory (eg, Traceable® Lollipop™ Waterproof/Shockproof Thermometer)

- Vortex mixer

- Interval timer

- Micropipette capable of dispensing 50 μL

- Alethia reader Meridian Bioscience, Inc. Catalog Number: 610189

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- Do not interchange Sample Preparation Apparatuses or Test Devices between lots.  Heat Treatment Tubes are interchangeable provided they are within assigned expiration date when used.

- Follow Biosafety Level 2 and Good Laboratory practices during testing.  Treat all specimens and used Test Devices as capable of transmitting infectious agents.  Do not eat, drink or smoke in areas where specimens or kit reagents are handled.

- Wear disposable gloves while handling specimens and thoroughly wash hands afterwards.

- Quality Control Programs for Molecular Testing Laboratories, including proper use and care of equipment, should be employed.

- The Alethia Group A Strep Test Device contains lyophilized reagents.  The protective pouch should not be opened until ready to perform the assay.

- The Alethia Group A Strep Test Device includes a latch feature that is designed to prevent contamination of the test area with amplification product.  Do NOT use Test Devices with broken latches.

- Dispose of used Alethia Test Devices immediately after processing, leaving the device latch securely in place.  Do NOT open the Test Device after processing.  Opening the device after amplification may result in contamination of the test area with amplification product.

STORAGE REQUIREMENTS:

  - The expiration date is indicated on the kit label.

  - Store the kit at 2-27 C.

At this facility, kits are stored:_______________________________________________

CALIBRATION:

Not applicable to this assay.

QUALITY CONTROL:

- This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

- Each device contains an internal control chamber that controls for amplification inhibition, assay reagents and sample processing effectiveness.

- The heat treatment step is monitored with an external thermometer and interval timer.  Use the max/min temperature memory of the thermometer to ensure that a temperature of 95 ± 5 C is maintained.  Use the interval timer to ensure that heat-treatment duration is 10 ± 2 minutes.

- Good laboratory practice recommends the use of control materials.  Users should follow the appropriate federal, state and local guidelines concerning the running of external quality controls.

- Alethia Group A Streptococcus External Control Reagents are supplied separately (Catalog 279910).  It is recommended that the reactivity of each new lot and each new shipment of Alethia Group A Strep be verified on receipt or before use.  External control tests should be performed thereafter in accordance with appropriate federal, state and local guidelines.  The Alethia Group A Strep test kit should not be used in patient testing if the external controls do not produce the correct results.

- A separate device must be used for each external control reagent.

QC Testing Frequency and Documentation:

For this facility, External QC is run: __________________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented: ___________________________________________________________________

- REAGENT PREPARATION

Ensure kit reagents are at room temperature (21-27 C) before use.  Incorrect results may be obtained if reagents are not brought to room temperature prior to use.

PROCEDURE:

- NOTE: Ensure that the Alethia reader is powered on and required performance verifications have been completed prior to initiation of SPECIMEN PREPARATION.  Refer to the Alethia incubator/reader Operator’s Manual for further information regarding instrument set-up and operation.

Specimen Preparation:

- 1.	Separate double swabs and remove one swab for Alethia testing.

2.	Add the throat swab to the Alethia SMP PREP.  Break off the handle of the swab.  Replace and secure SMP PREP cap.  Vortex the SMP PREP for a minimum of 10 seconds.  Specimens in the Alethia SMP PREP may be held at 2-29 C for up to 2 hours prior to heat treatment.

3.	Remove the tip cap from the SMP PREP and squeeze five to 10 drops of sample into a clean Alethia Heat Treatment Tube.

4.	Repeat Sample Preparation Steps for all samples to be processed.

5.	Heat each Sample/Control mixture in a dry-bath/heat block at 95 ± 5 C for 10 ± 2 minutes.  Monitor heat-treatment step with digital thermometer and interval timer.

6.	Remove each Heat Treatment Tube from the dry-bath/heat.  Heat-treated samples may be held at 21-29 C for up to one hour prior to testing.

TEST PROCEDURE:

- NOTE: A maximum of 10 samples can be processed in a single Alethia run.

- Vortex heat-treated samples for approximately 10 seconds.

- Remove 1 Alethia Group A Strep Test Device from its protective pouch per sample.  Carefully open the device, holding the chambers such that the lyophilized reagent will not fall out upon opening.  Place device on a flat surface or in a rack that can accommodate the device.

- Using a micropipette, first transfer 50 µL of the heat-treated sample to the TEST chamber (left/White Bead), and then transfer 50 µL of the heat treated sample to the CONTROL chamber (right/Yellow Bead) of the Alethia Test Device.  Take care not to introduce air to the reaction mixture.  Do not mix reactions with pipette.  Close the Alethia Test Device and fasten the latch securely.

- Tap device on the bench top or mix to remove air bubbles.  Carefully examine the test device for dissolution of the Control/Test Bead, for air bubbles left in the tube and liquid in the top of the device.  If undissolved beads, air bubbles or liquid is noted, tap the device on the bench top and repeat the visual inspection.

- Insert the Alethia Test Device into the Alethia reader and initiate amplification reaction and detection.  Results will be displayed at the conclusion of the run.

Alethia

CALCULATIONS:

There are no calculations associated with this procedure.

- INTERPRETATION OF RESULTS:

| Sample ID | Reported Result | Interpretation |
| --- | --- | --- |
| Patient Specimen | POSITIVE | Sample contains Streptococcus pyogenes target DNA. |
|  | NEGATIVE | No Streptococcus pyogenes DNA detected. |
|  | INVALID | No reportable result. Repeat the test using the original sample. Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure. |
| Positive Control | POSITIVE | Valid positive control result. Reagents active at time of use, Alethia performing correctly. |
|  | NEGATIVE | Incorrect control result. Repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | INVALID | No reportable result. Repeat entire assay run using original samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |
| Negative Control | POSITIVE | Incorrect control result. Repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor. |
|  | NEGATIVE | Valid negative control result. Reagents active at time of use, Alethia reader performing correctly. |
|  | INVALID | No reportable result. Repeat entire assay run using original samples. Improper sample preparation, reagent failure, instrument failure or internal control failure. |
| EMPTY WELL | NONE | No Alethia Test Device in the Alethia reader Well. OR The Alethia Test Device present is compromised due to sample preparation failure, dirty device or improperly seated device. Repeat the test using original sample. |

REPORTING OF RESULTS:

Positive Test: Group A Streptococcus (Streptococcus pyogenes) detected by NAAT or molecular method.

Negative Test: No Group A Streptococcus (Streptococcus pyogenes) detected.

This facility’s procedure for reporting patient’s results is: ________________________________ ______________________________________________________________________________

EXPECTED VALUES:

Overall incidence of Streptococcus pyogenes in patients tested during the 2012 clinical study was 14.5% (116/798).

ASSAY REACTIVITY:

The following S. pyogenes strains were tested and produced positive reactions at or below stated assay limit of detect of 400 CFU/Test: ATCC12384, NCIMB 13285, CCUG 33061, CCUG 33409, CCUG 39158, ATCC 49399, CCUG 53553.

Crossreactivity studies were performed with positive and negative specimens inoculated with bacterial or fungal organisms to a final concentration of 1.2 x 108 CFU/mL.  Human DNA was tested at 0.02 mg/mL with no cross-reactivity observed.  None of the following organisms reacted with Alethia Group A Strep:

Acinetobacter baumannii, Acinetobacter lwoffii, Aeromonas hydrophila, Arcanobacterium haemolyticum, Bordetella bronchiseptica, Bordetella holmesii, Bordetella parapertussis, Bordetella pertussis, Burkholderia cepacia, Campylobacter jejuni, Candida albicans, Citrobacter freundii, Corynebacterium diphtheria, Corynebacterium pseudodiphtheriticum, Enterococcus faecalis, Enterococcus faecium, Escherichia coli, Haemophilus influenza, Klebsiella oxytoca, Klebsiella pneumoniae, Lactococcus lactis, Legionella jordanis, Legionella micdadei, Legionella pneumophila, Listeria monocytogenes, Moraxella catarrhalis, Morganella morganii, Neisseria gonorrhoeae, Neisseria meningitidis, Proteus mirabilis, Proteus vulgaris, Pseudomonas aeruginosa, Pseudomonas fluorescens, Staphylococcus aureus, Staphylococcus epidermidis, Stenotrophomonas maltophilia, Streptococcus agalactiae, Streptococcus anginosus, Streptococcus bovis, Streptococcus canis, Streptococcus dysgalactiae (subspecies equisimilis), Streptococcus equinus, Streptococcus intermedius, Streptococcus mitis, Streptococcus mutans, Streptococcus pneumoniae, Streptococcus salivarius, Streptococcus suis, Streptococcus uberis, Streptococcus sp. Viridans type.

Two organisms produced unexpected results during initial testing that were not confirmed through further testing.  Arcanobacterium haemolyticum produced false-negative results for one of 13 replicates tested.  Haemophilus influenzae produced false-positive results for one of 13 replicates tested

Mycoplasma pneumoniae was tested at 1.5 x 106 CFU/mL with no reaction with the Alethia Group A Strep assay.

TESTS FOR INTERFERING SUBSTANCES:

Interfering substance testing was performed by adding potentially interfering substances diluted in sterile saline or contrived positive samples (ATCC 12344, ATCC 19615) to the Alethia SMP PREP.  Rayon swabs inoculated with negative throat flora and Liquid Amies Transport Media were added to the Alethia SMP PREP containing potentially interfering substances and tested.

The following biological substances, at the saturated solvent/diluent concentrations indicated, do not interfere with test results: Mucus (5.0mg/mL), Human saliva (10% v/v), and Whole Blood (2.5% v/v).  Whole Blood at concentrations greater than 2.5% v/v may interfere with the Alethia Group A Strep assay.

The following chemical substances, at the saturated solvent/diluents concentrations indicated, do not interfere with test results:

Acetaminophen (19.5 mg/mL), Aspirin (12.3 mg/mL), Cepacol® Mouthwash, [Cetyopridinium Chloride (0.005% v/v)], Cepacol® Sore Throat Lozenges [Benzocaine (0.09 mg/mL), Menthol (0.02 mg/mL)], Chloraseptic® Oral Anesthetic/Analgesic [Phenol (0.07% v/v)], Contac® Cold + Flu Tablets  [Acetaminophen (16.2 mg/mL),  Chlorpheniramine maleate (0.06 mg/mL), Phenylephrine HCl (0.16 mg/mL)], Crest® Complete Toothpaste [Sodium fluoride (0.1 mg/mL)], Diphenhydramine HCl (2.7 mg/mL), HALLS® Cough Drops [Menthol (0.08 mg/mL)], Ibuprofen (15.6 mg/mL), Listerine® Antiseptic Mouthwash [Eucalyptol (0.0092% v/v), Menthol (0.0042% v/v), Methyl salicylate (0.0060% v/v), Thymol (0.0064% v/v)], Robitussin® Cough/Chest Congestion Cough Syrup [Dextromethorphan HBr (0.2 mg/mL), Guaifenesin (2.0 mg/mL)].

Zicam® Oral Mist [Zincum Aceticum 2X, Zincum Gluonicum 1X (0.625% v/v)] produced invalid results in all replicates tested.

LIMITATIONS OF THE PROCEDURE:

- The Alethia Group A Strep assay does not distinguish between viable and non-viable organisms.

- Respiratory infections can be caused by Streptococci of serogroups other than A as well as other pathogens.  This device does not differentiate between carriers and infected individuals.

- The use of antibiotics or over-the-counter medications may suppress the Streptococcus growth in culture despite the presence of organisms detectable by nucleic acid tests.

- As will all diagnostic tests, results should be interpreted together with other information available to the physician.

- Zincum Aceticum 2X, Zincum Gluonicum 1X, as found in Zicam® Homeopathic Cold Remedies, interfere with the Alethia Group A Strep Assay.

PERFORMANCE CHARACTERISTICS:

- Refer to Directional Insert- Meridian Bioscience Alethia Group A Streptococcus

REFERENCES:

- Refer to Directional Insert- Meridian Bioscience Alethia Group A Streptococcus