[image: image1.png]meridian BIOSCIENCE
S





 Page 1 of 9

Procedure:  Alethia Mycoplasma Direct
Procedure and revision number: SN11039 CLSI 04/20
Meridian Catalog #480250

	Institution: 

	Address:

	Department:


	Prepared By:
	Date Adopted
	Supersedes procedure #:

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


	Distributed To:
	# of Copies
	Distributed To:
	# of Copies

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Alethia Mycoplasma Direct DNA amplification assay, performed on the Alethia Reader is a qualitative in vitro diagnostic test for the direct detection of DNA from Mycoplasma pneumoniae in human throat swabs obtained from patients suspected of having Mycoplasma pneumoniae infection. 

The Alethia Mycoplasma Direct assay utilizes loop-mediated isothermal DNA amplification (LAMP) technology to detect Mycoplasma pneumoniae by targeting a segment of the Mycoplasma pneumoniae genome. 

Results from the Alethia Mycoplasma Direct DNA amplification assay should be used in conjunction with clinical presentation, other laboratory findings, and epidemiological risk factors as an aid in the diagnosis of Mycoplasma infection and should not be used as the sole basis for treatment or other patient management.  Positive results do not rule out co-infection with other organisms and negative results in persons with respiratory tract infections may be due to pathogens not detected by this assay.  Lower respiratory tract infections due to M. pneumoniae may not be detected by this assay. If lower respiratory tract infection due to M. pneumoniae is suspected, additional laboratory testing using methods other than the Alethia Mycoplasma Direct DNA Amplification Assay may be necessary.

Mycoplasma pneumoniae is a common cause of human upper and lower respiratory infections, including pharyngitis, acute bronchitis and pneumonia.  Mycoplasma became recognized as a human pathogen in the 1960s, with Mycoplasma pneumoniae the best known and most studied.  M. pneumoniae contains a triple-layered cell membrane rather than a structured cell wall, which prevents identification by traditional gram stain methods and visualization by light microscopy. 

M. pneumoniae has been associated with up to 40% of community-acquired pneumonia cases. Infection occurs in both children and adults without geographical, gender or climate-related restrictions.  M. pneumoniae is most often associated with atypical pneumonia, presenting with symptoms that include headache, malaise, myalgia, fever, and sore throat accompanied by dry, paroxysmal cough.  While the clinical course of mycoplasmal pneumonia is generally mild and self-limiting, it has been associated with a mortality rate of approximately 1.4%.  An estimated 2 million cases of M. pneumoniae infection occur annually with approximately 100,000 pneumonia-related hospitalizations in the United States each year.  It is estimated as many as 18% of children requiring hospitalization with community-acquired pneumonia are caused by M. pneumoniae. 
Transmission of M. pneumoniae is generally from person to person by aerosol with a reported incubation period of one to four weeks. As patients with active infection carry Mycoplasma in the nose, throat, trachea, and sputum, the spread of the disease is facilitated by its accompanying cough.  This mode of transmission lends itself to outbreaks in close personal contact settings such as schools, military barracks, businesses, summer camps, colleges or institutions.  Similarly, the disease is commonly spread among family members within a household.  
Diagnosis of acute M. pneumoniae infection is difficult.  Basic diagnostic strategies in clinical practice include culture and serology.  M. pneumoniae culture is often impractical for patient management as the organism may take as many as six weeks to culture.  Serology products are widely available and offer practical diagnostic solutions, however, may be limited in their ability to detect acute infection.  Serological assays are often directed at specific IgM and IgG antibodies.  IgM antibodies are generally not detectable within the first seven days of symptom onset and may persist for months after active infection. Studies indicate that less than 50% of patients with acute infection demonstrate a positive IgG response. 

New diagnostic techniques, including DNA-amplification methods, may enable more rapid diagnosis and earlier identification of outbreaks for prevention of secondary cases through implementation of control measures.
SPECIMEN:
Preferred Sample Type: 

1. Throat swabs in Liquid Amies or Stuart’s transport containing sponge/pledget. 

2. Acceptable swabs:  Rayon, polyester, or flocked nylon with breakable plastic shafts. 

Undesirable Specimens: 
1. Throat swabs in Liquid Amies with charcoal transport media.  
2. Mini-tip swabs (absorbance capacity less than 60 µL) 
3. Eswabs
Collection and Preparation: Sample collection should be performed in accordance with institutional guidelines for collection of throat swab specimens.  Samples should be collected by vigorously swabbing the tonsils and the posterior pharynx. Swab sample(s) should be collected with suitable swab types (e.g. Rayon, Flocked Nylon, or Polyester).  

Place swab(s) in a non-nutritive transport medium (e.g. Liquid Amies, without charcoal; Liquid Stuart) containing pledget and transport to the laboratory.  Samples should be held at 2-29 C during transport.

Samples may be held at room temperature (19-29 C) for up to 24 hours prior to testing.  When testing will not be initiated within this time, the sample may be stored at 2-8 C for up to 14 days prior to testing.  

This facility’s procedure for specimen collection is: _____________________________________ ______________________________________________________________________________________
This facility’s procedure for transporting specimens is: __________________________________________ ______________________________________________________________________________________
This facility’s procedure for rejected specimens is: _____________________________________ ______________________________________________________________________________

MATERIALS AND EQUIPMENT:

MATERIALS PROVIDED:

The maximum number of tests obtained from this test kit is listed on the outer box.

1. Alethia Sample Preparation Apparatus II/Negative Control III (SMP PREP): Tris-buffered solution containing formalin-treated E. coli harboring S. aureus DNA and sodium azide (0.09%) as a preservative. 
2. Alethia Mycoplasma Test Device: Two-chambered device containing lyophilized amplification reagents (DNA polymerase, deoxynucleotide triphosphates) and either Mycoplasma-specific primers (TEST Chamber) or control primers (CONTROL Chamber)
3. Alethia Heat Treatment Tubes

MATERIALS NOT PROVIDED:

1. Alethia Mycoplasma Direct External Control Kit, Catalog Number: 479890

2. Disposable latex gloves, powder free

3. DNase/RNase-free, aerosol resistant pipette tips

4.    Specimen collection and transport materials:

· Swabs with breakable, plastic shaft (single or double swab): Polyester, Flocked Nylon, or Rayon

· Non-nutritive Transport Medium on Pledget/Sponge (Maximum Volume: 1.2 mL): Liquid Amies (without charcoal) or Liquid Stuart 

EQUIPMENT NOT PROVIDED:

1. Dry-bath with 12 mm heat block capable of 95 C

2. Digital thermometer with Max/Min Temperature Memory (eg, Traceable® Lollipop™ Waterproof/Shockproof Thermometer)

3. Vortex mixer

4. Interval timer

5. Micropipette capable of dispensing 50 μL

6. Alethia Reader, Meridian Bioscience, Inc. Catalog Number: 610189 
PERFORMANCE CONSIDERATIONS:

1. All reagents are for in vitro diagnostic use only.  
2. Mini-tip swabs (absorbance capacity less than 60 µL) are not acceptable for Alethia Mycoplasma Direct specimen collection.

3. Do not interchange the Sample Preparation Apparatus or Test Devices between lots.  Heat treatment tubes are interchangeable provided they are within assigned expiration dates when used.

4. Follow Biosafety Level 2 and Good Laboratory practices during testing.  Treat all specimens and used Test Devices as capable of transmitting infectious agents.  Do not eat, drink or smoke in areas where specimens or kit reagents are handled.

5. Wear disposable gloves while handling specimens and thoroughly wash hands afterwards. Change gloves often.

6. Quality Control Programs for Molecular Testing Laboratories should be employed. To avoid risk of false positives due to organism/amplicon contamination, employ prevention measures (eg, clean work area with suitable cleaner that destroys nucleic acid, use only aerosol resistant pipette tips, handle one specimen at a time) as required by Good Molecular Practices. 

7. The Alethia Mycoplasma Test Device contains lyophilized reagents.  The protective pouch should not be opened until ready to perform the assay.

8. The Alethia Mycoplasma Test Device includes a latch feature that is designed to prevent contamination of the test area with amplification product.  Do NOT use Test Devices with broken latches.

9. Dispose of used Alethia Test Devices immediately after processing, leaving the device latch securely in place.  Do NOT open the Test Device after processing. Opening the device after amplification may result in contamination of the test area with amplification product. 

STORAGE REQUIREMENTS:

The expiration date is indicated on the kit label.  Store the kit at 2-27 C.
At this facility, kits are stored: _______________________________________________

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting agencies.
1. Each device contains an internal control chamber that controls for amplification inhibition, assay reagents and sample processing effectiveness.

2. The heat treatment step is monitored with an external thermometer and interval timer.  Use the max/min temperature memory of the thermometer to ensure that a temperature of 95 ± 5 C is maintained.  Use the interval timer to ensure that heat-treatment duration is 10 ± 2 minutes.

3. Good laboratory practice recommends the use of control materials.  Users should follow the appropriate federal, state and local guidelines concerning the running of external quality controls.  

4. Alethia Mycoplasma Direct External Positive Control Reagent is supplied separately (Catalog 479890).  The Sample Preparation Apparatus II/Negative Control III reagent supplied within the Alethia Mycoplasma Direct kit (Catalog 480250) serves as the External Negative Control.  It is recommended that the reactivity of each new lot and each new shipment of Alethia Mycoplasma Direct be verified on receipt or before use.  External control tests should be performed thereafter in accordance with appropriate federal, state and local guidelines.  The Alethia Mycoplasma Direct test kit should not be used in patient testing if the external controls do not produce the correct results.

5. A separate device must be used for each external control reagent.

QC Testing Frequency and Documentation:

For this facility, External QC is run: __________________________________________________                                                                                                   

Results of External QC and action(s) taken when control results are unacceptable are documented: _______________________________________________________________________________
PROCEDURE:

NOTE: Ensure that the Alethia Reader is powered on and required performance verifications have been completed prior to initiation of SPECIMEN PREPARATION.  Refer to the Alethia Operator’s Manual for further information regarding instrument set-up and operation.

Specimen Preparation:

1. Add single throat swab to the Alethia SMP PREP.  Break off the handle of the swab.  Replace and secure SMP PREP cap.  Vortex the SMP PREP for a minimum of 10 seconds.  

2. Remove the tip cap from the SMP PREP and squeeze 5 - 10 drops of sample into a clean Alethia Heat Treatment Tube.

3. Repeat Sample Preparation Steps for all samples to be processed.

4. Heat each Sample/Control mixture in a dry-bath/heat block at 95 ± 5 C for 10 ± 2 minutes.  Monitor heat-treatment step with digital thermometer and interval timer.

5. Remove each Heat Treatment Tube from the dry-bath/heat block.  Heat-treated samples may be held at 19-29 C for up to 3 hours for repeat testing.  NOTE: Heat treatment must be re-performed (95 ± 5 C for 10 ± 2 minutes) according to Step 4 if heat-treated samples are stored prior to testing. 

TEST PROCEDURE:

NOTE: A maximum of 10 samples can be processed in a single Alethia Reader run. Refer to the Alethia Operator’s Manual for further information regarding instrument set-up and operation.
1. Vortex heat-treated samples for approximately 10 seconds. 

2. Remove 1 Alethia Mycoplasma Test Device from its protective pouch per sample.  Carefully open the device, holding the chambers such that the lyophilized reagent will not fall out upon opening.  Place device on a flat surface or in a rack that can accommodate the device.

3. Using a micropipette, first transfer 50 µL of the heat-treated sample to the TEST chamber (left/White Bead), and then transfer 50 µL of the heat-treated sample to the CONTROL chamber (right/Yellow Bead) of the Alethia Test Device.  Take care to not introduce air to the reaction mixture. Do not mix reactions with pipette.  Close the Alethia Test Device and fasten the latch securely.

4. Tap device on the bench top or mix to remove air bubbles.  Carefully examine the test device for dissolution of the Control/Test Bead, for air bubbles left in the tube and liquid in the top of the device.  If undissolved beads, air bubbles or liquid is noted, tap the device on the bench top and repeat the visual inspection.  Amplification and detection should be initiated within 15 minutes.
5. Insert the Alethia Test Device into the Alethia Reader and initiate amplification reaction and detection.  Results will be displayed at the conclusion of the run.
CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

	Sample ID 
	Reported Result
	Interpretation

	Patient Specimen
	POSITIVE
	Sample contains Mycoplasma pneumoniae target DNA.

	
	NEGATIVE
	No Mycoplasma pneumoniae DNA detected.

	
	INVALID
	No reportable result.  Repeat the test using the original sample.

Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure.

	Positive Control
	POSITIVE
	Valid positive control result.  Reagents active at time of use, Alethia Reader performing correctly.

	
	NEGATIVE
	Incorrect control result.  Patient results not reportable.  Repeat the control test as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor.

	
	INVALID
	No reportable result.  Repeat entire assay run using original samples. 

Improper sample preparation, reagent failure, instrument failure or internal control failure.

	Negative Control
	POSITIVE
	Incorrect control result. Patient results not reportable.  Repeat the control test as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor.

	
	NEGATIVE
	Valid negative control result.  Reagents active at time of use, Alethia Reader performing correctly.

	
	INVALID
	No reportable result.  Repeat entire assay run using original samples. 

Improper sample preparation, reagent failure, instrument failure or internal control failure.

	EMPTY WELL
	NONE
	No Alethia Test Device in the Alethia Reader Well.
OR

The Alethia Test Device present is compromised due to sample preparation failure, dirty device or improperly seated device.  Repeat the test using original sample.


REPORTING OF RESULTS:

Positive Test:  Mycoplasma pneumoniae DNA detected.
Negative Test: Mycoplasma pneumoniae DNA not detected or below the limit of detection.

This facility’s procedure for reporting patient’s results is: ________________________________ ______________________________________________________________________________

EXPECTED VALUES:

The incidence of Mycoplasma pneumoniae was established during clinical studies conducted in 2015-2016. The overall prevalence of M. pneumoniae in prospective throat swab specimens was 7.5% (34/456).  Prevalence by patient age is provided below. 

	Patient Age
	Total Positive
	Total Samples
	Prevalence (%)

	0 to 1 Month
	0
	7
	0.0%

	2 Months to 2 Years
	5
	157
	3.2%

	3 Years to 12 Years
	22
	147
	15.0%

	13 Years to 21 Years
	5
	38
	13.2%

	22 Years to 65 Years
	2
	86
	2.3%

	>65 Years
	0
	21
	0.0%


ASSAY REACTIVITY:

The following M. pneumoniae strains were tested with Alethia Mycoplasma and produced positive reactions at or below stated assay limit of detect of 2350 CFU/mL: 

PI 1428 (ATCC 29085), MAC (ATCC 15492), M52 (ATCC 15293), Bru (ATCC 15377), M129-B170 (ATCC 29343), Mutant 22 (ATCC 39505) UAB 55612, UAB 56317, UMTB-10G (ATCC 49899).  Testing demonstrated that both Type 1 and Type 2 strains are detected with the assay.  The same reactivity is expected with the Alethia Mycoplasma Direct assay.

Crossreactivity studies were previously performed with the Alethia Mycoplasma DNA Amplification assay.  Positive and negative specimens were inoculated with potentially interfering bacterial or fungal organisms to minimum final concentrations of 1.0 x 106 CFU/mL.  Viruses were tested at concentrations greater than 1.0 x 105 TCID50/mL or 1.0 x 106 copies/mL.  Human DNA was tested at 2.0 ng/test with no crossreactivity observed.  Positive specimens contained M. pneumoniae concentrations near the limit of detection.  None of the following organisms or materials were identified as crossreactive or interfering in the Alethia Mycoplasma assay:

Acinetobacter baumannii, Acinetobacter calcoaceticus, Actinomyces odontolyticus, Bacillus subtilis, Bacteroides fragilis, Bordetella parapertussis, Bordetella pertussis, Burkholderia cepacia, Candida albicans, Candida glabrata, Candida parapsilosis, Chlamydia pneumoniae, Citrobacter freundii, Clostridium difficile, Corynebacterium diphtheriae, Enterobacter cloacae, Enterococcus faecalis, Escherichia coli, Escherichia coli (ESBL), Fusobacterium nucleatum, Haemophilus ducreyi, Haemophilus influenzae, Haemophilus parainfluenzae, Helicobacter pylori, Klebsiella pneumoniae, Klebsiella pneumoniae (KPC), Legionella pneumophila, Listeria monocytogenes, Moraxella catarrhalis, Mycoplasma genitalium, Mycoplasma hominis, Neisseria cinerea, Neisseria gonorrhoeae, Neisseria meningitidis, Nocardia asteroides, Peptostreptococcus anaerobius, Proteus mirabilis, Proteus vulgaris, Pseudomonas aeruginosa, Salmonella paratyphi (Group A), Salmonella typhimurium (Group B), Serratia liquefaciens, Staphylococcus aureus, Staphylococcus epidermidis, Streptococcus agalactiae (Group B), Streptococcus anginosus (Group F), Streptococcus bovis (Group D), Streptococcus canis (Group G), Streptococcus equisimilis, Streptococcus mitis, Streptococcus pneumoniae, Streptococcus pyogenes, Streptococcus salivarius, Ureaplasma urealyticum, Adenovirus, Coronavirus, Coxsackievirus, Cytomegalovirus, Epstein Barr virus, Herpes simplex virus 1, Herpes simplex virus 2, Human metapneumovirus, Influenza A, Influenza B, Parainfluenza virus 1, Parainfluenza virus 2, Parainfluenza virus 3, Respiratory syncytial virus A, Respiratory syncytial virus B, Rhinovirus, Human DNA.  No crossreactivity is expected with the Alethia Mycoplasma Direct assay.
TESTS FOR INTERFERING SUBSTANCES:

Potentially interfering substances were tested with simulated negative and contrived low positive (M. pneumoniae strains M129 and FH) samples.  The following substances, at the saturated solvent/diluent concentrations indicated, do not interfere with Alethia Mycoplasma Direct assay: 

Acetaminophen (18.1 mg/mL), Albuterol Sulfate (20 mg/mL), Aspirin (9.1 mg/mL), Azithromycin dehydrate (2.0 mg/mL), Cepacol®  Mouthwash [Alcohol (1.4% v/v), Cetylpyridinium chloride (0.005% v/v)], Contac® Cold + Flu Tablets [Acetaminophen (14.8 mg/mL), Chlorpheniramine maleate (0.06 mg/mL), Phenylephrine HCl (0.15 mg/mL)], Diphenhydramine HCl (2.6 mg/mL), Erythromycin (20.0 mg/mL), HALLS® Cough Drops [Menthol (0.06 mg/mL)], Ibuprofen (12.7 mg/mL), Phenylephrine HCl (0.595 mg/mL), Prednisone (20.0 mg/mL), *Robitussin® Cough+Chest Congestion Cough Syrup [Dextromethorphan HBr (0.20 mg/mL), Guaifenesin (2.0 mg/mL)], Saline Nasal Spray [Sodium chloride (0.65 mg/mL)], Mucus (5.0mg/mL), and White Blood Cells (0.5% v/v). 

Whole blood was found to produce invalid results at concentrations greater than 2%.  

Phenylephrine HCl at concentrations greater than 0.595 mg/mL may produce false negative results with low positive samples.

* Robitussin® Cough+Congestion Cough Syrup initially produced 1/3 invalid results with Positive Sample P2 that produced acceptable results during repeat testing (1/1 positive).  Robitussin® Cough+Congestion Cough Syrup is not considered an interferent.

LIMITATIONS OF THE PROCEDURE:

1. This product can be used only with the Alethia instrument.

2. The Alethia Mycoplasma Direct DNA assay is a qualitative assay and does not provide quantitative values or information about organism load.

3. This device has not been evaluated for monitoring treatment of Mycoplasma pneumoniae infections.

4. Respiratory infections can be caused by Mycoplasma pneumoniae as well as other pathogens. Positive results do not preclude coinfection with other respiratory pathogens. 

5. Performance of Alethia Mycoplasma Direct has been established with prospective human throat swab specimens only. Performance has not been established for other specimen types (nasopharyngeal specimens, samples from immunocompromised individuals, or from asymptomatic patients.)

6. The detection of nucleic acid is dependent upon proper specimen collection, handling, transportation, storage and preparation.  Failure to observe proper procedure in any one of these steps can lead to incorrect results. 

7. Prevalence of M. pneumoniae infection will affect the test’s predictive value.

8. Organism nucleic acids may persist in vivo independent of organism viability. Detection of analyte DNA target may not imply causative agents for clinical symptoms.  The Alethia Mycoplasma Direct assay does not differentiate between carriers and infected individuals.

9. Results should be interpreted together with clinical history, epidemiological data, and other information available to the physician.  As with all molecular based tests, (A) False negative results may occur from the presence of inhibitors, technical error, sample mix-up or low numbers of organisms in the clinical specimen, (B) False positive results may occur from the presence of cross-contamination by target organisms, their nucleic acids or amplified product, and from non-specific signals.
10. Nasal decongestants containing Phenylephrine HCl at concentrations greater than 0.595 mg/mL may produce false negative results with low positive specimens.
11. Whole blood at concentrations greater than 2% may interfere with the Alethia Mycoplasma Direct assay.
12. It has been observed that Starplex collection devices can generate a higher rate of invalid results when used in conjunction with the Alethia Mycoplasma Direct assay.
PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert- Meridian Bioscience Alethia Mycoplasma Direct
REFERENCES:

Refer to Directional Insert- Meridian Bioscience Alethia Mycoplasma Direct
Meridian Bioscience

3471 River Hills Drive

Cincinnati, OH 45244 USA



Ph: (800) 343-3858, (513) 271-3700

SN11039 REV 04/20

