Procedure:  Premier Platinum HpSA PLUS

Procedure and revision number: SN11203 CLSI Rev 04/20

Meridian Catalog #: 601396

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Premier Platinum HpSA PLUS enzyme immunoassay (EIA) is an in vitro qualitative procedure for the detection of Helicobacter pylori antigens in human stool.  Test results are intended to aid in the diagnosis of H. pylori infection and to monitor response during and post-therapy in patients.  Accepted medical practice recommends that testing by any current method, to confirm eradication, be done at least four weeks following completion of therapy.

The Premier Platinum HpSA PLUS test utilizes a plurality of monoclonal anti-H. pylori capture antibodies adsorbed to microwells. (Plurality is defined as a mixture of monoclonal antibodies.)  Diluted patient samples and a conjugate (peroxidase conjugated to a plurality of monoclonal antibodies) are added to the wells and incubated for one hour at room temperature.  A wash is performed to remove unbound material.  Substrate is added and incubated for 10 minutes at room temperature.  Color develops in the presence of bound enzyme.  Stop Solution I is added and the results are interpreted visually or spectrophotometrically.

SPECIMEN:

Preferred sample type and volumes:

1.  Liquid or semi-solid stool:  200-300 uL (100 uL for testing and the remaining for retesting as needed)

2.  Formed solid stool:  20 mm sample (5-6 mm for testing and the remaining for retesting if   needed)

Undesirable specimen characteristics: Stool in transport media, swabs, or preservatives are inappropriate for testing.

Collection and Storage:

The specimen should be received in an airtight transport container and stored at 2-8 C until tested.  The specimen should be tested as soon as possible but may be held up to 72 hours at 2-8 C prior to testing.  (See SPECIMEN PREPARATION section for instructions on diluting samples.)  If testing cannot be performed within this time frame, specimens should be frozen immediately upon receipt and stored frozen (-20 C to –80 C) until tested.  Specimens may be frozen and thawed twice.

NOTE: Stool in transport media, swabs, or preservatives are inappropriate for testing.

This facility’s procedure for specimen collection is: _________________________________ ___________________________________________________________________________

This facility’s procedure for transporting specimens is: ______________________________

__________________________________________________________________________

This facility’s procedure for rejected specimens is: _________________________________ __________________________________________________________________________

Specimen preparation:

      - Using a pipetting device, add 500 µL of Sample Diluent to a clean test tube.

- Mix stool as thoroughly as possible prior to pipeting.

- Liquid or semi-solid stools - Using the supplied transfer pipette, add 100 µL (second mark from the tip of the pipette) of stool into Sample Diluent.  Using same pipette, gently withdraw and expel the stool suspension several times, then vortex 15 seconds.  Save the transfer pipette in the sample for later use.

- Formed/Solid stools - Using a wooden applicator stick, transfer a small portion (5-6 mm diameter) of thoroughly mixed stool into Sample Diluent.  Emulsify stool using the wooden applicator stick, then vortex 15 seconds.

- Stool specimens may be centrifuged after dilution.  Centrifuge at approximately 2750 x G for five minutes or until solid matter separates from liquid.  Proceed with the assay after recovering supernate.

This facility’s procedure for sample labeling is:_____________________________________ ___________________________________________________________________________

MATERIALS AND EQUIPMENT:

MATERIALS PROVIDED:

The maximum number of tests obtained from this test kit is listed on the outer box.

- Antibody Coated Microwells - Breakaway plastic wells coated with a plurality of murine monoclonal antibodies specific for H. pylori.

- Positive Control - Inactivated H. pylori diluted in 10 mM phosphate-buffered solution with 0.02% Thimerosal, pH 7.2.

- Sample Diluent/Negative Control - pH 7.2, 10 mM phosphate-buffered solution with 0.02% Thimerosal.

- Premier 20X Wash Buffer I - pH 6.8, 180 mM phosphate-buffered solution with 0.2% Thimerosal.

- Enzyme Conjugate – A plurality of murine monoclonal antibodies specific for H. pylori conjugated to horseradish peroxidase in a pH 7.8, 50 mM Tris-buffered solution containing 0.02% Thimerosal.

- Premier Substrate I - a-buffered solution containing urea peroxide and tetramethylbenzidine. (pH 5.0)

- Premier Stop Solution I – 1 M phosphoric acid. CAUTION: Avoid contact with skin. Flush with water if contact occurs.

- Transfer pipettes (one for each test sample).  Each pipette is marked to indicate 50 µL, 100 µL, 200 µL and 300 µL volumes.

- Plate sealer.

- Wooden stick applicators.

MATERIALS NOT PROVIDED:

- Test tubes (12 x 75 mm) for dilution of sample

- Distilled or deionized water

- Squirt bottle

- Graduated cylinder for making 1X Wash Buffer I

EQUIPMENT:

- EIA plate reader capable of reading absorbance at 450 or 450/630 nm*

- Semiautomated plate washer (e.g., BioTek Elx50) *

*Note: It is the operator’s responsibility to validate the semiautomated plate washers and readers prior to their use with this product.

REAGENT PREPARATION:

- Bring the entire kit, including microwell pouch, to 19-27 C before use.

- Prepare 1X Wash Buffer I as needed. For example: 4.0 mL of Premier 20X Wash Buffer + 76.0 mL of distilled or deionized water is sufficient to wash one strip.  Place in a clean squirt bottle.  The 1X Wash Buffer I can be stored at 19-27 C for up to three months.

Performance Considerations:

- All reagents are for in vitro diagnostic use only.

- Patient specimens may contain infectious agents and should be handled and disposed of as potential biohazards.

- All reagents should be mixed gently before use.

- Do not interchange the Microwells, Enzyme Conjugate, Substrate I Reagent, or Positive Control reagents between lots.  (The Sample Diluent, Premier 20X Wash Buffer I and Premier Stop Solution I are interchangeable provided the reagents are within their assigned expiration dates when used.)

- Allow reagents to warm to 19-27 C before use.

- Hold reagent vials vertically at suitable distance above the well to insure proper drop size and delivery.

- Do not use kit components beyond labeled expiration date.

- Replace colored caps on correct vials.

- Dispose of used wash buffer and all test materials in an appropriate container.  Treat waste as a potential biohazard.

- The Positive Control reagent contains inactivated H. pylori.  It should be handled, however, as a potential biohazard.

- Avoid skin contact with Premier Stop Solution I (1 M phosphoric acid).  Flush with water immediately if contact occurs.

- Do not reuse microwells.

- Unused microwells must be placed back inside resealable pouch.  It is important to protect strips from moisture.

- The transfer pipettes provided with this kit must be used for specimen preparation and transfer.  Use one per specimen.

- Avoid splashing when dispensing diluted stool into microwells by placing the transfer pipette tip about halfway down the well and dispensing slowly down the side of well.

- Microwell washing is to be performed precisely as directed in assay procedure.  Inadequate washing may be the cause of elevated background in any EIA protocol.

- All reagents except the Premier 20X Wash Buffer I are provided already diluted to the proper concentration.

- Any deviation below or above set incubation times may affect sensitivity and specificity and should be avoided.

- Stool must be mixed thoroughly (regardless of consistency) to insure a representative sample prior to pipetting.

- Some precipitation may occur in Premier 20X Wash Buffer I when it is stored at 2-8 C.  The precipitate will dissolve when a working dilution is made with the Wash Buffer.

- Do not use vials that lack a label, a lot number, or an expiration date.

SHELF LIFE AND STORAGE:

The expiration date is indicated on the kit label. Store all kit reagents as 2-8 C and return the kit promptly to the refrigerator after each use.

At this facility, kits are stored: _______________________________________________

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

- At the time of each use, kit components should be visually examined for obvious signs of microbial contamination, freezing or leakage.  Do not use contaminated or suspect reagents.  The Positive and Negative Controls must be used with each test batch.  See the INTERPRETATION OF RESULTS section for a description of the expected results for control reagents.  Tests should be considered invalid when either control reagent does not produce the expected results.  In such cases, repeat tests and controls.  If, on repeat testing, the expected reactions are still not observed, and the reagents are still within their expiration date, contact Meridian’s Technical Support Services at 1-800-343-3858 for assistance (USA) or your local distributor.

- The controls are used to monitor reagent reactivity.  Failure of the controls to produce the expected results can mean that one or more of the reagents are no longer reactive at the time of use, the test was not performed correctly, or that reagents or samples were not added.  The Positive Control will not ensure precision at the cut-off.

- Suspect failure of the washing method or device if the Negative Control and/or Positive Controls consistently produce out of specification results.  Increasing the number of washes, washing more vigorously, decanting more thoroughly or recalibrating washing devices should correct the problem.  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

- Specimen matrix interference has not been observed in this assay as samples are significantly diluted before testing in Sample Diluent.  For this reason, the positive and negative control reagents supplied as part of this assay are prepared in the matrix of the Sample Diluent.  If control materials that are identical in composition to test samples are preferred, the user can prepare these by diluting known positive and negative specimens in Sample Diluent according to the SPECIMEN PREPARATION section of this insert.  Add 100 µL of user prepared controls to test wells.

QC Testing Frequency and Documentation:

For this facility, External QC is run: ______________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented: _______________________________________________________________

PROCEDURE:

  - After the pouch has reached temperature, break off the required number of microwells (1 well for each specimen, plus 1 positive and 1 negative control well per batch).  Place the microwells in the microwell strip holder and record the location of all wells.  Unused microwells must be resealed in the pouch immediately.

  - Using the specimen transfer pipette, add 100 µL of diluted stool (second mark from the tip of the pipette) to the appropriate well.  (Place the pipette tip halfway into well and let the sample slowly run down side of well.)

  - Add 2 free falling drops of Positive Control and 100 µL of Sample Diluent /Negative Control to the appropriate wells.

  - Add 1 free falling drop (approximately 50 µL) of Enzyme Conjugate to each well.  Firmly shake/swirl the plate for 30 seconds.

  - Cut plate sealer to size and press firmly onto top of microwells to seal.  Incubate the plate for 1 hour at 19-27 C.

  - Carefully remove the plate sealer and wash wells:

    - Manual method:

      - Dump plate contents firmly into a biohazard receptacle.

      - Bang the inverted plate on a clean stack of paper towels.

      - Fill all wells with 1X Wash Buffer I, directing stream of buffer to the sides of the wells to avoid foaming.

      - Repeat wash cycle (dump, bang on fresh towels, fill) 4 times for a total of 5 wash cycles.  After the last fill, dump and bang plates on fresh towels hard enough to remove as much excess wash buffer as possible, but do not allow wells to completely dry at any time.

    - Semiautomated method using validated equipment

      - Aspirate the contents of the well.

      - Fill the wells to the top (approx 300-350 µL/well) with 1X Wash Buffer I then aspirate.  The washer manifold should be adjusted to ensure no foaming occurs during the filling of the wells and that the wells are thoroughly aspirated after each wash.

      - Repeat step ii a minimum of 4 more times.  Following the last wash, test wells should be thoroughly aspirated to remove as much moisture as possible.

  - Clean the underside of all wells with a lint-free tissue.

  - Add 2 free falling drops (approx. 100 µL) of Premier Substrate Solution I to each well.  Firmly shake/swirl the plate for 30 seconds.  Incubate for 10 minutes at 19-27 C.

  - Add 2 free-falling drops (approx. 100 µL) of Premier Stop Solution I to each well.  Firmly shake/swirl the plate for 30 seconds.

Note: Initial color of positive reaction is blue, which changes to yellow upon addition of Premier Stop Solution I.

  - Inspect and record reactions.  Test results can be read visually or using a spectrophotometric reader.

    - Visual Determination - Read within 15 minutes after adding Premier Stop Solution I.

    - Spectrophotometric Determination - Zero EIA reader on air.  Wipe underside of wells with a lint-free tissue.  Read absorbance at 450 nm or 450/630 nm within 15 minutes of adding Premier Stop Solution I.

CALCULATIONS:

There are no calculations associated with the procedure.

INTERPRETATION OF RESULTS:

The following interpretations apply to both initial diagnosis and monitoring of anti-H. pylori therapy.

Visual Reading

Negative = colorless to faint yellow

Positive   = definite yellow color

To be called positive, a faint yellow color must be confirmed by a spectrophotometric reading.  If a spectrophotometer is not available, the cut-off must be determined by an alternative method.

Spectrophotometric Single Wavelength (450 nm)

Negative:  < 0.140

Positive:  ( 0.140

Negative Control: < 0.140

Positive Control: ( 0.640

Spectrophotometric Dual Wavelength (450/630 nm)

Negative: < 0.100

Positive:  ( 0.100

Negative Control: < 0.100

Positive Control: ( 0.600

If a Negative Control is < 0.000, reblank the plate reader to air and reread the plate.

Extremely strong positive reactions may yield a purple precipitate within a few minutes of stopping the reaction.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:

A positive result indicates the presence of H. pylori antigens.  A negative result indicates the absence of H. pylori antigens, or that the level of antigens is below what can be detected by the assay.    The magnitude of the OD above the cut-off is not indicative of the severity or extent of H. pylori infection, nor can it be correlated to an endpoint titer.  Extremely strong positive reactions may yield a purple precipitate within a few minutes of stopping the reaction.

| Result Obtained | Suggested Report |
| --- | --- |
| Positive | H. pylori antigen present |
| Negative | H. pylori antigen not present or below the level of detect |
| Invalid | DO NOT REPORT ANY RESULT. REPEAT TEST. |

This facility’s procedure for patient result reporting is:__________________________________ ____________________________________________________________________________________________________________________________________________________________

EXPECTED VALUES:

Studies on the epidemiology of H. pylori have shown that this organism is present worldwide.  Gastritis caused by H. pylori has been shown to correlate with age, ethnic background, family size and socioeconomic class. The prevalence of H. pylori infection in a given population can vary from 20% to 90%.  In patients diagnosed with duodenal ulcers, however, it has been shown in every age group to be approximately 80%. Currently recommended eradication treatments have an efficacy rate between 75% and 90%.

The Premier Platinum HpSA PLUS test detects the presence of H. pylori antigens in human stool.  Expected values for a given population should be determined for each laboratory. The rate of positivity may vary depending on geographic location, method of specimen collection, handling and transportation, test employed and general health environment of patient population under study.  As demonstrated by Premier Platinum HpSA in tests conducted in the United States, Canada and Italy, incidence of disease ranged from 34% to 53% to 69% respectively.

TESTS FOR INTERFERING SUBSTANCES:

The following substances, that may be present in human stool, do not interfere with positive or negative test results at the stated concentrations per 500 µL human stool:

Barium sulfate – 25 mg, Mylanta 11.5 mg, Pepto Bismol  – 0.44 mg, Prilosec (omeprazole) – 1 mg, Tagamet (cimetidine) – 1 mg, Tums  – 10 mg, Hemoglobin – 62.5 mg,   Mucin – 17 mg, NSAID Ibuprofen – 0.25 mg, Stearic acid – 5.3 mg, Palmitic acid – 2.65 mg, White blood cells –   250 µL, Whole blood – 250 µL

LIMITATIONS OF THE PROCEDURE:

1. 	The test is qualitative and no quantitative interpretation should be made with respect to the values.

2. 	Test results should be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

3. 	Antimicrobials, proton pump inhibitors and bismuth preparations are known to suppress H. pylori and ingestion of these prior to H. pylori testing (culture, histology, rapid urease, UBT, antigen) may give a false negative result.  If a negative result is obtained for a patient ingesting these compounds within two weeks prior to performing the Premier Platinum

HpSA PLUS test, it may be a false-negative result and the test should be repeated on a new specimen obtained two weeks after discontinuing treatment.  A positive result for a patient ingesting these compounds, within two weeks prior to performing the Premier Platinum HpSA PLUS test, should be considered accurate.  As an example, patients with H. pylori were placed on a proton pump inhibitor (Lansoprazole) or bismuth for two weeks, and tested with the Premier Platinum HpSA and a urea breath test.  Patients were then taken off treatment for two weeks and retested.  At the end of treatment, both assays were negative for some patients, but returned to positive 2 weeks post-treatment (see table).

|  |  | Premier Platinum HpSA | Breath Test |
| --- | --- | --- | --- |
| Treatment | Time Point | Pos/Total | % Positive | Pos/Total | % Positive |
| PPI | End of Treatment | 15/20 | 75.0% | 12/20 | 60.0% |
|  | 2 Weeks Post-Treatment | 19/20 | 95.0% | 18/20 | 90.0% |
| Bismuth | End of Treatment | 15/20 | 75.0% | 11/20 | 55.0% |
|  | 2 Weeks Post-Treatment | 19/20 | 95.0% | 18/20 | 90.0% |

4. 	Performance characteristics have not been established for watery, diarrheal stools.

5. 	Performance characteristics have not been established in asymptomatic populations.

6. 	H2 blockers do not interfere with positive results.

PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience Premier Platinum HpSA PLUS

REFERENCES:

Refer to Directional Insert – Meridian Bioscience Premier Platinum HpSA PLUS