Procedure:  premier® WB Lyme IgG

Procedure and revision number: CS612030  Rev. 03/19

Meridian Catalog:  612030

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might effect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Premier WB Lyme IgG test kit is intended for the qualitative determination of immunoglobulin class IgG antibodies against Borrelia burgdorferi in human serum and plasma (K+-EDTA, Li+-heparin, Na+-citrate) samples, that have been found positive or borderline/equivocal using an enzyme immunoassay (EIA) or immunofluorescence assay (IFA) test procedure for B. burgdorferi antibodies.  Results can be read manually or automated utilizing EUROLineScan.  This test is used as an aid in the diagnosis of infections with B. burgdorferi and the associated diseases, in conjunction with other laboratory and clinical findings.

The test kit contains test strips with electrophoretically separated antigen extracts of B. burgdorferi.  The blot strips will be blocked and incubated in the first reaction step with diluted patient samples.  In the case of positive samples, specific antibodies of the class IgG (and IgA, IgM) will bind to the antigens.  To detect the bound antibodies, a second incubation is carried out using an enzyme-labelled anti-human IgG (enzyme conjugate), catalyzing a color reaction.

SPECIMEN:

Preferred sample types:

Human serum or plasma (K+-EDTA, Li+-heparin, Na+-citrate) obtained by venipuncture

Undesirable specimens:

Hemolyzed, hyperlipemic, hemolytic, heat-treated or contaminated samples

Collection and Storage:

Human serum or plasma (K+-EDTA, Li+-heparin, Na+-citrate).  Samples are to be obtained by venipuncture.

Samples should be stored at room temperature no longer than 8 hours.  If the assay will not be completed within 8 hours, the samples should be refrigerated at +2 C to +8 C.  If the assay will not be completed within 48 hours, or if the samples will be stored beyond 48 hours, samples should be frozen at -20 C or lower.  Samples should not be repeatedly frozen and thawed.  Frozen samples must be mixed well after thawing and prior to testing.  Diluted samples should be incubated within 8 hours.  Do not use bacterially contaminated samples.  It is the responsibility of the individual laboratory to use all available references and/or its own studies to determine its own specific stability criteria.

MATERIALS AND EQUIPMENT:

Materials Provided:

| Component | Format | Symbol |
| --- | --- | --- |
| Test strips strips with electrophoretically separated B. burgdorferi antigens | 30 x 1 |  |
| Evaluation matrix with control strip 1 test strip incubated with a positive control serum | 1 sheet | --- |
| Enzyme conjugate alkaline-phosphatase labelled anti-human IgG (goat), 10x concentrate | 2 x 3 mL |  |
| Universal buffer 10x concentrate | 1 x 100 mL |  |
| Substrate solution nitroblue tetrazolium chloride/5-bromo-4-chloro-3-indolylphosphate (NBT/BCIP), ready for use | 1 x 50 mL |  |

Materials Not Provided:

- Pipettes with a range of 10 µL to 1000 µL with single-use tips

- Tweezers (for handling of the strips)

- Multistep pipette with top for dispensing up to 1.5 mL per pass

- Rocker

- Vortex mixer

- Distilled or deionized water for buffer preparation, stopping

- Glass/plastic tubes for dilution or samples and reagents

- Lint-free towel

- Filter paper

- Black tray (Please note, each channel within the tray is for single use only)

- For Visual Evaluation or use with use with automated sample processors, order tray only (catalog # 11768)

- For Automated Evaluation, the tray is included in Automated Evaluation Accessory Pack (catalog # 610216A)

For visual evaluation of the strips, please use the required Premier Western Blot Reaction Control Card (catalog # 11764) and the Premier Western Blot Lyme IgG Evaluation Form (available on the Meridian Bioscience, Inc. website).

For automated evaluation of the strips with EUROLineScan only:

- For automation purposes, it is necessary to order the Premier Western Blot Automated Evaluation Accessory Pack (catalog # 610216A).  This pack contains the following items:

- Green paper (5 sheets)

- Adhesive foil for approx 30 test strips (2 foils)

- Black tray for 30 strips (1 tray).  Please note, each channel within the tray is for single use only

- For the creation of work protocols and the evaluation of test strips using EUROLineScan:

- Catalog # 610216: EUROLineScan scanner

- Catalog # 610214: EUROLineScan software

PERFORMANCE CONSIDERATIONS:

- For in vitro diagnostic use.  For use by laboratory professionals in a clinical or research laboratory setting.

- Before starting the assay, carefully read the instructions.  Use only the valid version provided with the kit.

- Do not substitute or mix with reagents from other manufacturers.

- Observe prudent laboratory practice and safety guidelines.  Avoid eye and skin contact with specimens and reagents.  In case of eye or skin contact, wash off thoroughly with plenty of clean running water.  Remove and wash contaminated clothing.  In case of ingestion, obtain medical attention.

- The controls of human origin have tested negative for HBsAg, anti-HCV, anti-HIV-1 and anti-HIV-2.  Nonetheless, all materials should be treated as being a potential infection hazard and should be decontaminated or disposed of with proper biohazard precautions.  Some of the reagents contain sodium azide in a non-declarable concentration as an antibacterial agent.  Avoid skin contact.

- Waste disposal: Patient samples and incubated blot strips should be handled as infectious waste.  Follow local, state and federal regulations regarding handling and disposal of hazardous waste, including potentially explosive materials such as accumulated or concentrated sodium azide.

STORAGE REQUIRMENTS:

Storage and Stability: The test kit has to be stored at a temperature between +2 C to +8 C.  Do not freeze.  Unopened, all test kit components are stable until the indicated expiry date.  All reagents must be brought to room temperature (+18 C to +25 C) approx. 30 minutes before use.  Unopened, the reagents are stable until the indicated expiry date when stored at +2 C to +8 C.  After initial opening, the reagents are stable until the indicated expiry date, if stored at +2 C to +8 C and protected from contamination, unless stated otherwise.

At this facility, kits are stored:_______________________________________________

CALIBRATION:

Not applicable to this assay.

TEST PROCEDURE:

Reagent Preparation:

- Coated test strips: Ready for use.  Do not open until the test strips have reached room temperature to prevent condensation on the strips.  After removal of the strips the packing should be sealed tightly and stored at +2 C to +8 C.  To ensure correct evaluation of results, the strip lot number must match the number on the evaluation matrix.

- Enzyme conjugate: The enzyme conjugate is supplied as a 10x concentrate.  For the preparation of the working-strength enzyme conjugate the amount required should be removed from the bottle using a clean pipette and diluted 1:10 with working-strength universal buffer.

For example: For 1 test strip dilute 0.15 mL anti-human IgG concentrate with 1.35 mL of the working-strength universal buffer.

The working-strength enzyme conjugate should be used at the same working day.

- Universal buffer: The universal buffer is supplied as a 10x concentrate.  For the preparation of the working-strength universal buffer the amount required should be removed from the bottle using a clean pipette and diluted 1:10 with deionized or distilled water.

For example: For the incubation of 1 test strip 1.5 mL buffer concentrate should be diluted with 13.5 mL deionized or distilled water.

The working-strength buffer should be used at the same working day.

- Substrate solution: Ready for use.  Close the bottle immediately after use, as the contents are sensitive to light .

Specimen Preparation:

The patient samples to be investigated are diluted 1:51 in ready for use diluted universal buffer.

For example: Add 30 µL of sample to 1.5 mL ready for use diluted universal buffer and mix well by vortexing (sample pipettes are not suitable for mixing).

This facility’s procedure for specimen preparation is:____________________________  _____________________________________________________________________

This facility’s procedure for sample labeling is:_________________________________ ______________________________________________________________________

Test Procedure:

- Blocking:

  - According to the number of serum samples to be tested fill each channel of the incubation tray with 1.5 mL ready for use diluted universal buffer.

  - Remove the required amount of blot strips from the packing using a pair of tweezers.  Insert the blot strip in the incubation tray such that the number on the test strip is visible.

  - Incubate for 15 minutes at room temperature (+18 C to +25 C) on a rocking shaker.  Afterwards aspirate off all the liquid.

- Sample Incubation (1st step):

  - Fill each channel with 1.5 mL of the diluted serum samples.  One serum sample per channel.

  - Incubate at room temperature (+18 C to +25 C) for 30 minutes on a rocking shaker.

- Washing:

  - Aspirate off the liquid from each channel and wash 3 x 5 minutes each with

1.5 mL working-strength universal buffer on a rocking shaker.

4.  Conjugate-incubation (2nd step):

- Pipette 1.5 mL ready for use diluted enzyme conjugate (alkaline phosphatase conjugated anti-human IgG) into each channel.

- Incubate for 30 minutes at room temperature (+18 C to +25 C) on a rocking shaker.

5. Washing:

- Aspirate off the liquid from each channel.  Wash as described above.

6. Substrate-incubation (3rd step):

- Pipette 1.5 mL substrate solution into the channels of the incubation tray.

- Incubate for 10 minutes at room temperature (+18 C to +25 C) on a rocking shaker.

7.  Stopping:

- Aspirate off the liquid from each channel and wash each strip 3 x 1 minute with deionized water.

8.  Evaluation:

Results should be read once strips have dried. Results can be read either by manual/visual interpretation or automated using EUROLineScan. It is recommended to evaluate the strips at the day of incubation. See Evaluation of Results section.

## Premier WB Lyme IgG Incubation Protocol

RESULTS:

EVALUATION OF RESULTS:

- After stopping the reaction using deionized or distilled water, apply adhesive foil to the Lyme IgG Evaluation Form.  (Note:  for Automated Evaluation, print an evaluation protocol on green paper using the EUROLineScan software; refer to EUROLineScan user manual).  Place the incubated blot strips onto the corresponding field of the Evaluation Form using a pair of tweezers.  The position of the blot strips can be corrected while they are wet.  As soon as all blot strips have been placed onto the Evaluation Form, they should be blotted using filter paper to remove excess moisture.  After they have dried for 1 hour, the blot strips will be fixed to the adhesive foil.

- The test strip lot-specific evaluation matrix included in the kit is fitted with a blot strip originating from the same gel as the other blot strips in the test kit.  This blot strip has been incubated with a positive control serum and indicates the exact position of the relevant bands.  The blot strips are labelled with a specific number.  This number refers to the strip lot and is also printed on the corresponding evaluation matrix.  The strip lot number on the evaluation matrix must match the strip lot number on the blot strip.

- For evaluation, put the evaluation matrix against the blot strip and align the black band above the strip lot number with the alignment bar of the evaluation matrix.

Antigens on the test strips:

| Band | Antigen | CDC Relevant Bands |
| --- | --- | --- |
| 83/93 kDa | Membrane-vesical protein, p 83/93 degradation product of p 100 | X |
| 66 kDa | p 66 | X |
| 60 kDa | p 60 | - |
| 58 kDa | p 58 | X |
| 45 kDa | p 45 | X |
| 41 kDa | Flagellin, p 41 | X |
| 39 kDa | Borrelia membrane protein A, BmpA, p 39 | X |
| 34 kDa | Outer surface protein B, OspB, p 34 | - |
| 31 kDa | Outer surface protein A, OspA, p 31 | - |
| 30 kDa | p 30 | X |
| 28 kDa | p 28 | X |
| 25 kDa | Outer surface protein C, OspC, p 25 | X |
| 18/21 kDa | p 18/21 | X |

Visual evaluation with Reaction Control Card:

Hold the reaction control card (catalog # 11764) next to the incubated strip and compare the band intensity with the two color bars of the card.  Band intensities weaker than the positive color bar are evaluated as negative.  Band intensities stronger than the positive color bar are evaluated as positive.  Band intensities stronger than the negative color bar and weaker than the positive color bar are evaluated as borderline.

Note: When blot strips are evaluated visually with the Reaction Control Card, staining could occasionally completely or partially mask the bands making them unreadable.  Masking may appear as heavy back-ground staining, heavy speckling or blotching.  Due to variations in test performance and the uncertainty associated with unreadable Western blots, it is recommended that all unreadable blots be retested, using the original specimen.  If the original specimen repeatedly yields unreadable blots, and symptoms persist, a fresh specimen should be tested in 2 to 4 weeks.  By utilizing EUROLineScan for automated evaluation of the strips, the software calculates and considers the background intensity to avoid misinterpretation.

Automated evaluation with EUROLineScan:

The evaluation protocol (printed on green paper) with the test strips is scanned using a flatbed scanner (catalog # 610216) and evaluated with EUROLineScan software (catalog # 610214).  For use of the EUROLineScan program please refer to the EUROLineScan user manual.  The code for entering the test into EUROLineScan is B.b._WB_US_G.

QUALITY CONTROL:

Each blot strip contains a control band.  This must show an intense staining after the incubation, in order for the results to be valid.

Additional controls may be tested in accordance with local, state and/or federal regulations or accreditation requirements and your laboratory’s quality control procedures.  It is recommended that the user refer to CLSI document C24-A and 42 CFR 493.1256 for guidance on appropriate quality control practices.

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________

Results of External QC and action(s) taken when control results are unacceptable are

documented:________________________________________________________________

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Band cut-off determination:

The cut-off of the Premier WB Lyme IgG test system is defined as the lowest limit of a clearly visible band.  Since visual examination by the operator might be subjective, the reaction control card was developed to standardize strip evaluation.  Alternatively, the EUROLineScan software was established to allow for automated evaluation.

IgG class antibodies against B. burgdorferi:

Positive:  According to the CDC criteria, the IgG blot is considered as positive if any 5 of the following 10 bands are positive: p18/21, p25 (OspC) p28, p30, p39 (BmpA), p41, p45, p58, p66 and p83/93.

Negative:  A negative interpretation denotes that IgG antibodies against less than 5 of the 10 significant B. burgdorferi proteins were detected, or no IgG antibodies against significant B. burgdorferi proteins were detected.  If a negative result is obtained and infection with B. burgdorferi is clinically suspected, a new sample should be tested 3 to 4 weeks later.

Note: Borderline results are considered as negative.  It is known that B. burgdorferi-infected individuals who have recently seroconverted may display incomplete patterns but may develop increased reactivity (both in intensity and number) when followed for a period of four to six months.

In the event this test becomes inoperable, this facility’s course of action for patient samples

is:_______________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Report test results as, “Negative or below the level of detect for Borrelia burgdorferi IgG antibodies.”

Positive test:  Report test result as, “Positive for Borrelia burgdorferi IgG antibodies.”

EXPECTED VALUES:

The range of values and positivity of different populations among different studies are presented below with available patient demographics.

| Population | n | Sex | Mean age & range | Premier WB Lyme IgG |
| --- | --- | --- | --- | --- |
|  |  |  |  | p18/21 | p25 | p28 | p30 | p39 | p41 | p45 | p58 | p66 | p83/93 | Anti-B. burgd. positive (%) |
| Prospective study |
| EIA positive | 304 | 169 men, 133 women, 2 unknown | 56 yrs 4 - 88 yrs 4 unknown | 239 | 197 | 111 | 130 | 251 | 297 | 184 | 190 | 176 | 190 | 233 (76.6%) |
| Lyme disease |
| Sensitivity study | 101 | 35 men, 58 women, 8 unknown | 45 yrs 16 - 87 yrs 9 unknown | 47 | 33 | 5 | 31 | 49 | 96 | 27 | 28 | 49 | 28 | 56 (55.4%) |
| Normal population study |
| Endemic | 98 | 89 men, 9 women | 34 yrs 18 - 56 yrs | 10 | 3 | 1 | 7 | 4 | 84 | 6 | 14 | 14 | 10 | 3 (3.1%) |
| Non-endemic | 100 | 82 men, 18 women | 38 yrs 3 - 62 yrs | 9 | 2 | 0 | 13 | 2 | 87 | 4 | 5 | 16 | 8 | 0 (0.0%) |

Note: It is recommended that each laboratory determine its own normal range based on the population and the equipment used.

CROSSREACTIVITY:

Cross-reactivity was investigated using characterized samples from the following groups as shown in the tabl below.  All the samples tested were found negative with the Premier WB Lyme IgG according to the CDC criteria.

| Panel | n | Premier WB Lyme IgG |
| --- | --- | --- |
|  |  | Negative (%) |
| EBV | 15 | 15 (100.0%) |
| HSV | 14 | 14 (100.0%) |
| Influenza viruses | 15 | 15 (100.0%) |
| Helicobacter pylori | 11 | 11 (100.0%) |
| Measles | 15 | 15 (100.0%) |
| Parvovirus B19 | 12 | 12 (100.0%) |
| Rubella | 14 | 14 (100.0%) |
| Treponema | 10 | 10 (100.0%) |
| CMV | 10 | 10 (100.0%) |
| Babesiosis | 3 | 3 (100.0%) |
| Anaplasmosis (ehrlichiosis) | 10 | 10 (100.0%) |
| Rickettsial diseases | 4 | 4 (100.0%) |
| Rheumatoid arthritis | 22 | 22 (100.0%) |
| Total | 155 | 155 (100.0%) |

Note: The results obtained with babesiosis (3) and rickettsial diseases (4) samples are not conclusive as the number of tested samples was insufficient.

INTERFERING SUBSTANCES:

Hemolytic, lipemic and icteric samples showed no influence at the result up to a concentration of 500 mg/dl for hemoglobin, 2000 mg/dl for triglycerides and 40 mg/dl for bilirubin in this test system.  Interferences with albumin, intralipids and cholesterol have not been investigated.

LIMITATIONS OF THE PROCEDURE:

- This kit is used as an aid in diagnosis only.  A positive result should be interpreted together with clinical findings and other serological tests.  The results obtained from this assay are not diagnostic proof of the presence or absence of a disease.

- Negative results do not exclude the possibility of infection with B. burgdorferi.

- Interference from albumin, intralipids and cholesterol with the assay performance has not been investigated.

- Insufficient washing (e.g. too small wash buffer volumes, or shortened reaction times) can lead to incorrect results.

- The test is only to be performed by laboratory professionals in a clinical laboratory setting.  Strictly adhere to the test procedure.  Deviations from the indicated volumes, times etc. can lead to incorrect results.

- Adaptation of this assay for use with automated sample processors and other liquid handling devices, in whole or in part, may yield in differences in test results from those obtained using the manual procedure.  It is the responsibility of each laboratory to validate that its automated procedure yields test results within acceptable limits.

## PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience premier® WB Lyme IgG

## REFERENCES:

Refer to Directional Insert – Meridian Bioscience premier® WB Lyme IgG

# Text boxes

15 min

Blocking:

Put the blot strip and 1.5 mL working-strength

universal buffer into the incubation channel

Wash:

Aspirate off, wash 3 x 5 min with 1.5 mL working-strength universal bufferPlus each

2. Incubation with conjugate:

Aspirate off, pipette 1.5 mL of working-strength

enzyme conjugate (1:10) into the incubation channel

3. Incubation with substrate:

Aspirate off, pipette 1.5 mL of substrate solution into the incubation channel

Stopping:

Aspirate off and wash

30 min

10 min

Evaluation:

Evaluate the same day

##### Shake

##### Shake

##### Shake

Wash:

Aspirate off, wash 3 x 5 min with 1.5 mL working-strength universal bufferPlus each

1. Sample incubation:

Aspirate off, pipette 1.5 mL of diluted serum sample (1:51)

into the incubation channel

30 min

##### Shake