Procedure:  Premier CAMPY

Procedure and revision number: SN11174 CLSI 07/20

Meridian Catalog #: 618096

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

Premier CAMPY enzyme immunoassay (EIA) is an in vitro qualitative procedure for the detection of specific Campylobacter antigens in stool samples from patients with signs and symptoms of gastroenteritis. Premier CAMPY detects C. jejuni and C. coli in human stool that may be either unpreserved or preserved in Cary-Blair-based transport media. Test results are to be used in conjunction with information obtained from the patient’s clinical evaluation and other diagnostic procedures.

Premier CAMPY is not intended for point-of-care use. The device is intended for use in hospital, reference, regional, private or state laboratory settings.

Premier Campy is an enzyme immunoassay for the direct detection of Campylobacter antigens in human stool samples.  Breakaway microwells are coated with Campylobacter-specific monoclonal antibodies.  Diluted patient specimen is added to the microwells and incubated.  Upon completion of the incubation, a wash step is performed to remove unbound material and a Horseradish Peroxidase (HRP)–Anti–Campylobacter conjugate is added to the washed microwells.  If Campylobacter antigens are present, an antibody-enzyme complex is formed.  A second wash step is performed to remove unbound materials and a chromagen substrate is added to the microwells.  A blue color develops in the presence of bound enzyme.  Premier Stop Solution I is added, changing the initial blue reaction to yellow.  Test results are interpreted visually or spectrophotometrically.

SPECIMEN:

Preferred sample types:  Unpreserved stool, or stool placed in a Cary-Blair-based transport media

Undesirable specimens:  Stool in non-Cary-Blair-based transport media

Collection and Storage:

Human stool samples, unpreserved: Samples should be received in an airtight transport container and stored at 2–8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 96 hours at 2–8 C.  (See SPECIMEN PREPARATION section for instructions on diluting samples.)  Samples that will not be tested within 96 hours should be frozen immediately upon receipt and stored at ≤ –20 C until tested.  Specimens may be frozen and thawed twice.

Human stool samples, preserved in Cary-Blair-based media: Samples should be stored at 2–8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 96 hours at 2–8 C.  (See SPECIMEN PREPARATION section for instructions on diluting samples.)  Samples that will not be tested within 96 hours should be frozen immediately upon receipt and stored at ≤ –20 C until tested.  Specimens may be frozen and thawed twice.

This facility’s procedure for specimen collection is: _________________________________ ___________________________________________________________________________

This facility’s procedure for transporting specimens is: ______________________________

__________________________________________________________________________

This facility’s procedure for rejected specimens is: _________________________________ __________________________________________________________________________

This facility’s procedure for sample labeling is:_____________________________________ ___________________________________________________________________________

MATERIALS AND EQUIPMENT:

MATERIALS PROVIDED:

- Premier Campy Microwells: Monoclonal antibody–coated microwells.  The antibodies are specific to C. jejuni and C. coli.

- Premier Campy Enzyme Conjugate: HRP-conjugated monoclonal antibodies specific to C. jejuni and C. coli in a buffered protein solution containing 0.1% ProClin® and 0.03% gentamicin as preservatives.

- Premier 20X Wash Buffer III: Concentrated wash buffer containing 0.1% ProClin® and 0.3% gentamicin as preservatives.

- Premier Substrate I: Buffered solution containing urea peroxide and tetramethylbenzidine at pH 5.0.

- Premier Stop Solution I: 1M Phosphoric acid.

- Premier Campy Sample Diluent/Negative Control: Buffered protein solution containing 0.1% ProClin® and 0.03% gentamicin as preservatives.

- Premier Campy Positive Control: Inactivated C. jejuni in a buffered protein solution containing 0.094% sodium azide and 0.03% gentamicin as preservatives.

- Transfer pipettes

- Microwell strip holder

- Microwell plate sealers

MATERIALS NOT PROVIDED:

- Disposable latex gloves

- Test tubes (eg, 10 x 75 mm or 12 x 75 mm) for dilution of sample

- Distilled or deionized water

- Squirt bottle

- Graduated cylinder for making 1X Wash Buffer

- Absorbent paper

- Wooden applicator sticks

- Waste container with disinfectant and/or autoclavable biohazard bags

- Vortex mixer

- Interval timer

- EIA microplate reader capable of reading absorbance at 450 nm or 450/630 nm* (optional)

- StatFax™-2200 Incubator/Shaker (StatFax™ is a trademark of Awareness Technology, Inc.)* (optional)

- Semi-automated microplate washer* (optional)

*Note: It is the operator’s responsibility to validate StatFax™, semi-automated plate   washers and readers prior to their use with this product.

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- Directions should be read and followed carefully.

- Do not interchange the Microwells, Enzyme Conjugate, Premier Substrate I or Positive Control reagents between lots.  (The Sample Diluent/Negative Control, Premier 20X Wash Buffer III and Premier Stop Solution I are interchangeable provided the reagents are within their assigned expiration dates when used.)

- Do not use kit components beyond labeled expiration date.

- Do not use vials that lack a label, a lot number, or an expiration date.

- Do not use any reagent if it is discolored or turbid.  Discoloration or turbidity may be a sign of microbial contamination.

- Allow reagents to warm to 21–27 C before use.

- All reagents should be mixed gently before use.

- Hold reagent vials vertically at suitable distance above the well to insure proper drop size and delivery.

- Replace colored caps on correct vials.

- Do not reuse microwells.

- Unused microwells must be placed back inside resealable pouch.  It is important to protect strips from moisture.

- The transfer pipettes provided with this kit must be used for specimen preparation and transfer.  Use one per specimen.

- Avoid splashing when dispensing diluted stool into microwells by placing the transfer pipette tip about halfway down the well and dispensing slowly down the side of well.

- Microwell washing is to be performed precisely as directed in the assay procedure.  Inadequate washing may be the cause of elevated background in any EIA protocol.

- All reagents except the Premier 20X Wash Buffer III are provided already diluted to the proper concentration.

- Any deviation below or above set incubation times may affect sensitivity and specificity and should be avoided.

- Stool must be mixed thoroughly (regardless of consistency) to ensure a representative sample prior to pipetting.

WARNINGS

- Patient specimens may contain infectious agents and should be handled and disposed of as potential biohazards.

- Dispose of used wash buffer and all test materials in an appropriate container.  Treat waste as a potential biohazard.

- The Positive Control reagent contains inactivated C. jejuni.  It should be handled, however, as a potential biohazard.

- Avoid skin contact with Premier Stop Solution I.  Flush with water immediately if contact occurs.

SHELF LIFE AND STORAGE:

The expiration date is indicated on the kit label.  Store the kit at 2-8 C and return the kit promptly to the refrigerator after each use.

At this facility, kits are stored:_______________________________________________

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

- At the time of each use, kit components should be visually examined for obvious signs of  microbial contamination, freezing or leakage.  Do not use contaminated or suspect reagents.

- The performance of Premier CAMPY should be verified using the Positive Control and Sample Diluent/Negative Control with each test batch.  See the INTERPRETATION OF RESULTS section above for a description of the expected results for control reagents.  Tests should be considered invalid when either control reagent does not produce the expected results.  In such cases, repeat tests and controls.  If, on repeat testing, the expected reactions are still not observed with in-date reagents, call Meridian’s Technical Support Services at 800-343-3858 (USA) or your authorized Meridian distributor for assistance.

- The controls are used to monitor reagent reactivity.  Failure of the controls to produce the expected results can mean that one or more of the reagents are no longer reactive at the time of use, that the test was not performed correctly, or that reagents or samples were not added.  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or  your local distributor.

- Specimen matrix interference has not been observed in this assay as samples are significantly diluted before testing in Sample Diluent.  For this reason, the Positive and Negative Controls supplied as part of this assay are prepared in matrices similar to the Sample Diluent.  If control materials that are identical in composition to test samples are preferred, the user can prepare these by diluting known positive and negative specimens in Sample Diluent according to the SPECIMEN PREPARATION section of this insert.  Add 100 µL of user prepared controls to test wells.

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented:________________________________________________________________

REAGENT PREPARATION:

            - Bring the entire kit, including microwell pouch, to 21-27 C before use.

2.	Prepare 1X Wash Buffer as needed. For example: 10.0 mL of Premier 20X Wash Buffer III + 190.0 mL of distilled or deionized water is sufficient to wash one strip.  The 1X Wash Buffer can be stored at 2-27 C for up to one month.  Bring to 21-27 C prior to use.

SPECIMEN PREPARATION:

PROCEDURAL NOTES:

The Premier CAMPY transfer pipette is diagrammed below.

Mix stool as thoroughly as possible prior to pipetting.

1.	Human stool samples, unpreserved:

  - With the dropper assembly (or equivalent), add 200 µL of Sample Diluent/Negative Control to a small test tube.

  - Formed/Solid stools:

    - Use a wooden applicator stick (or equivalent) to add a small portion (3–4 mm diameter) of thoroughly mixed stool to the Sample Diluent/Negative Control test tube.

    - Emulsify the stool using the wooden applicator stick.

    - Vortex the emulsion for a minimum of 15 seconds.

  - Liquid/Semi-Solid stools:

    - Using a transfer pipette provided with the kit, add 50 µL (first mark from pipette tip) of thoroughly mixed stool to the Sample Diluent/Negative Control test tube.

    - Vortex the mixture for a minimum of 15 seconds.

    - Save the transfer pipette in the test tube for future use.

2.	Human stool samples, preserved in Cary-Blair-based media:

  - With the dropper assembly (or equivalent), add 200 µL of Sample Diluent/Negative Control to a small test tube.

  - Using a transfer pipette provided with the kit, add 200 µL (third mark from pipette tip) of thoroughly mixed stool to the Sample Diluent/Negative Control test tube.

  - Vortex the mixture for a minimum of 15 seconds.

  - Save the transfer pipette in the test tube for future use.

TEST PROCEDURE:

- After the pouch has reached temperature (21–27 C), break off the required number of microwells (1 well for each specimen, plus 1 positive and 1 negative control well per batch).  Place the microwells in the microwell strip holder and record the location of all wells.  Unused microwells must be resealed in the pouch immediately.

- Using the specimen transfer pipette, add 100 µL of diluted stool (second mark from the tip of the pipette) to the appropriate well.  (Place the pipette tip halfway into well and let the sample slowly run down side of well.)

- Add 2 free falling drops of Positive Control.  With a transfer pipette, add 100 µL (second mark from the tip of the pipette) of Sample Diluent/Negative Control to the appropriate wells.  Cut the plate sealer to size and press it firmly onto the top of the microwells to seal.

- Incubate the plate for 1 hour at 21-27 C.  Alternatively, laboratories equipped with a shaker/incubator (StatFax™-2200) can incubate and rotate the plate for 30 minutes at 24 C and at Setting 5 (1000 rpm).

- Carefully remove the plate sealer and wash the wells:

        - Manual method:

- Dump the plate contents firmly into a biohazard receptacle.

- Invert and bang the plate on a clean stack of paper towels.

- Fill all wells with 1X Wash Buffer, directing a stream of buffer to the sides of the wells to avoid foaming.

- Repeat the wash cycle (dump, bang on fresh towels, fill) 4 more times for a total of 5 wash cycles.  After the last fill, dump and bang plate on fresh towels hard enough to remove as much excess wash buffer as possible, but do not allow wells to completely dry at any time.

b.	Semi-automated wash method:

- Dump the plate contents firmly into a biohazard receptacle.

- Invert and bang the plate on a clean stack of paper towels, then place empty plate on washer device.

- Fill the wells to the top (approx. 300-350 µL/well) with 1X Wash Buffer and then aspirate.  The washer manifold should be adjusted to ensure no foaming occurs during the filling of the wells and that the wells are thoroughly aspirated after each wash.

- Repeat step iii 4 more times.  Following the last wash, test wells should be thoroughly aspirated to remove as much moisture as possible, but do not allow wells to completely dry at any time.

- Add 2 free falling drops (approximately 100 µL) of Enzyme Conjugate to each well.  Firmly shake/swirl the plate for 30 seconds.

- Cut the plate sealer to size and press it firmly onto the top of the microwells to seal.  Incubate the plate for 30 minutes at 21–27 C.  Alternatively, laboratories equipped with a shaker/incubator (StatFax™-2200) can incubate and rotate the plate for 15 minutes at 24 C and at Setting 5 (1000 rpm).

- Carefully remove the plate sealer and wash the wells as described in Step 5.

- Add 2 free-falling drops (approximately 100 µL) of Premier Substrate I Solution to each well.  Firmly shake/swirl the plate for 30 seconds.  Incubate the plate for 10 minutes at 21–27 C.

- Add 2 free-falling drops (approximately 100 µL) of Premier Stop Solution I to each well.  Firmly shake/swirl the plate for 30 seconds.  Note: The initial color of a positive reaction is blue, which changes to yellow upon addition of Premier Stop Solution I.

- Clean the underside of all wells with a lint-free tissue.

- Inspect and record reactions.  Test results can be read visually or using a spectrophotometric reader.

- Visual Determination - Read within 15 minutes after adding Premier Stop Solution I.

- Spectrophotometric Determination – Zero EIA reader on air.  Read absorbance at 450 nm or 450/630 nm within 15 minutes of adding Premier Stop Solution I.

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Visual Reading

Negative = Colorless to very faint yellow

Positive   = Definite yellow color

A very faint yellow color must be evaluated by a spectrophotometric reading.

Spectrophotometric Single Wavelength (450 nm)

Negative:  < 0.150

Positive:    ≥ 0.150

Negative Control:  < 0.150

Positive Control:    ≥ 0.600

Spectrophotometric Dual Wavelength (450/630 nm)

Negative:  < 0.100

Positive:    ≥ 0.100

Negative Control:  < 0.100

Positive Control:    ≥ 0.600

If a Negative Control is < 0.000, re-blank the plate reader to air and reread the plate.

A positive result indicates the presence of Campylobacter antigens. A negative result indicates the absence of Campylobacter antigens, or that the level of antigens is below what can be detected by the assay. The magnitude of the OD above the cutoff is neither indicative of the severity or extent of Campylobacter infection, nor can it be correlated to an endpoint titer. Extremely strong positive samples may yield either an intense yellow color or a purple precipitate within a few minutes of stopping the reaction. In this case, the spectrophotometer may yield an “out” reading. This reading is considered a positive.

If the frequency of low positive results (OD between 0.150 and 0.200 for single wavelength

and between 0.100 and 0.150 with dual wavelength) is greater than 5% of specimens tested, this may indicate insufficient washing. More vigorous washing or increasing the washes to seven washes in step 5 of the procedure is recommended.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Report test results as “Campylobacter antigen not detected or the level of antigens is below detection of assay”.

Positive test: Report test results as “Campylobacter antigen detected”.

This facility’s procedure for patient result reporting is:__________________________________ ____________________________________________________________________________________________________________________________________________________________

EXPECTED VALUES:

The performance of Premier CAMPY was evaluated during 2008 in several geographic regions of the United States.  The incidence of positive samples (either as C. jejuni or C. coli) during this period was approximately 1.5%.  The expected frequency for an individual laboratory may differ from this number since it is dependent on factors such as locale, the time of the year, and whether an outbreak has occurred.

ASSAY SPECIFICITY/CROSSREACTIVITY:

Crossreactivity studies were performed with positive and negative stool specimens inoculated with bacterial or fungal organisms to a final concentration of 1.1 x 108 CFU/mL or virus ranging from 1.3 x 104 to 3.1 x 106 TCID50/mL.  None of the following organisms in stool reacted with Premier CAMPY:

Aeromonas hydrophila, Bacteroides fragilis, Campylobacter fetus, Campylobacter lari, Candida albicans, Clostridium difficile, Clostridium perfringens, Enterobacter cloacae, Enterococcus faecalis, Escherichia coli, Escherichia coli O157:H7, Escherichia fergusonii, Escherichia hermannii, Helicobacter pylori, Klebsiella pneumoniae, Lactococcus lactis, Listeria monocytogenes, Peptostreptococcus anaerobius, Plesiomonas shigelloides, Proteus vulgaris, Pseudomonas aeruginosa, Pseudomonas fluorescens, Salmonella Groups B-E, Serratia liquefaciens, Serratia marcescens, Shigella boydii, Shigella dysenteriae, Shigella flexneri, Shigella sonnei, Staphylococcus aureus, Staphylococcus epidermidis, Vibrio parahaemolyticus, Yersinia enterocolitica, Adenovirus Types 40 and 41, Coxsackievirus, Echovirus, Rotavirus.

TESTS FOR INTERFERING SUBSTANCES:

The following substances, at the specified saturated solvent/diluent concentrations, do not interfere with test results in the final concentrations listed: Barium sulfate (5 mg/mL), fecal fat (equivalent to 2.65 mg stearic plus 1.3 mg palmitic acids per mL), hemoglobin (as

methhemoglobin) (3.2 mg/mL), Imodium AD® (0.00667 mg/mL), Kaopectate® (0.87mg/mL), mucin (3.33 mg/mL), Mylanta® (4.2 mg/mL), Pepto-Bismol® (0.87 mg/mL), Prilosec® (0.5 mg/mL), Tagamet® (0.5 mg/mL), TUMS® (0.5 mg/mL), whole blood (5%v/v).

LIMITATIONS OF THE PROCEDURE:

- The test is qualitative, and no quantitative interpretation should be made with respect to the values.

- Premier CAMPY cannot be used as the sole means of determining Campylobacteriosis.  Test results must be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

- Overincubation of the tests may lead to an increase in false-positive test results.  Conversely, incubation for periods less than those defined in this insert can result in an increase in false-negative tests.  Follow incubation times defined in this insert.

- Suspect failure of the washing method or device if the Negative and/or Positive Controls consistently produce out-of-specification results.  Increasing the number of washes, washing more vigorously, decanting more thoroughly or recalibrating washing devices should correct the problem.  If problems persist, contact Meridian’s Technical Support Services at 800-343-3858 (USA) or your authorized Meridian distributor for assistance.

PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience Premier CAMPY

REFERENCES:

Refer to Directional Insert – Meridian Bioscience Premier CAMPY