Procedure:  ImmunoCard®  Toxins A & B

Procedure and revision number: SN11223 CLSI 05/20

Meridian Catalog #712050

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.  Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

ImmunoCard Toxins A & B is a rapid, qualitative, horizontal-flow enzyme immunoassay (EIA) for detecting Clostridium difficile Toxins A and B in human stool. This assay is used as an aid in the diagnosis of C. difficile-associated disease.

In the US, ImmunoCard Toxins A & B is not intended for point-of-care use. The device is intended for moderately complex laboratories. In Canada, this device is not intended for point-of-care use.

ImmunoCard Toxins A & B consists of a membrane held in a plastic frame with two sample ports and two reaction ports.  The membrane carries immobilized antibodies to toxins A and B.  The Enzyme Conjugate Reagent consists of antibodies to toxins A and B coupled to horseradish peroxidase.  To perform the test, patient stool sample is diluted with Sample Diluent and Enzyme Conjugate and the mixture is incubated for five minutes to 24 hours.  During the incubation, molecules of toxin, if present, are bound to the anti-toxin antibodies of the conjugate.  Following incubation, an aliquot of the mixture is added to each of the two sample ports and the test is incubated for an additional five minutes at 20-26 C.  During the second incubation the toxin-conjugate complex is separated from particulate matter as the fluid portion of the sample flows through the membrane to the TEST and CONTROL reaction ports.  The toxin-conjugate complexes are then captured at the TEST reaction port by immobilized antitoxin in the reaction membrane.  (The second of the two reaction ports serves as an internal control.)  Both reaction ports are subsequently washed with ImmunoCard Wash Buffer I to reduce interference by contaminating proteins before ImmunoCard Substrate I is added.  The reaction ports are incubated for an additional five minutes during which time the Enzyme Conjugate modifies the ImmunoCard Substrate I.  The result is the appearance of a blue color.  Reactions are read visually.  Development of a blue color in the TEST reaction port indicates a positive test.  In the CONTROL port, the anti-toxin antibodies of the conjugate bind directly to the immobilized toxin.  The appearance of blue in the CONTROL reaction port indicates that sample was added, that reagents were active at the time of use and that proper sample migration occurred.

SPECIMEN:

Preferred sample types:

Unpreserved Stool

Undesirable specimens:

Stool samples in preservative

Collection:

The stool sample should be transported in an airtight container and stored at 2-8 C until tested.  The sample should be tested as soon as possible, but may be held up to four days at 2-8 C prior to testing.  If testing cannot be performed within this time frame, samples should be frozen immediately on receipt and stored frozen   (( -20 C) until tested.  Samples that are frozen and thawed up to three times are acceptable for use.

This facility’s procedure for specimen collection is: _________________________________ ___________________________________________________________________________

This facility’s procedure for transporting specimens is: ______________________________

__________________________________________________________________________

This facility’s procedure for rejected specimens is: _________________________________ __________________________________________________________________________

MATERIALS AND EQUIPMENT:

MATERIALS PROVIDED:

The maximum number of tests obtained from this kit is listed on the outer box.

- ImmunoCard Toxins A & B Test Card: A membrane pad housed in a plastic frame and enclosed in a foil pouch with a desiccant. The pad carries immobilized monoclonal anti-toxin A and goat polyclonal anti-toxin B at the TEST reaction port and C. difficile toxin at the CONTROL reaction port. Store the test cards at 2-8 C when not in use.

- Sample Diluent*: A buffered protein solution containing gentamicin and thimerosal (0.02%) as a preservative. Store at 2-8 C when not in use.

- Positive Control*: Inactivated crude C. difficile toxin in a buffered solution containing thimerosal (0.02%) as a preservative. The reagent is supplied ready for use. Store at 2-8 C when not in use.

- Enzyme Conjugate*: A blend of goat polyclonal antibodies to toxins A and B conjugated to horseradish peroxidase and suspended in a buffered protein solution containing gentamicin and thimerosal (0.02%). Store at 2-8 C when not in use.

- ImmunoCard Wash Buffer I*: A buffered solution containing thimerosal (0.01%) as a preservative. Store at 2-8 C when not in use.

- ImmunoCard Substrate I*: A buffered solution containing tetramethyl-benzidine and peroxide. Store at 2-8 C when not in use.

7. Plastic transfer pipettes with measuring marks for 25 μL and 150 μL.

* In a plastic dropper vial. Use as supplied.

MATERIALS NOT PROVIDED:

- Disposable latex gloves that should be used during the handling of the fecal samples as they are considered potentially biologically hazardous material

- Vortex for suspending of the stool sample in the Sample Diluent

- Interval timer

- Applicator sticks

- Small test tubes, such as 10 x 75 mm or 12 x 75 mm

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- Do not store at temperatures above 2-8 C.  Do not freeze.

- Do not deviate from the established insert method or falsely positive or falsely negative results may occur.

- Some patient samples contain infectious agents; therefore all patient samples should be handled and disposed of as if they are biologically hazardous.

- DO NOT interchange the Test Card, Enzyme Conjugate and Positive Control reagents from different kit lot numbers.  The other reagents (Sample Diluent, ImmunoCard Wash Buffer I, ImmunoCard Substrate I) can be interchanged between kits providing the reagents are within their expiration periods.  Do not use any reagent beyond its labeled expiration date.

- All stool samples must be mixed thoroughly before testing, regardless of consistency, to ensure a representative sample prior to testing.

- Failure to bring samples and reagents to room temperature (20 - 26 C) before testing may decrease assay sensitivity.

- Inspect Test Cards before removing the foil pouch.  Do not use Test Cards that have holes in the foil pouch or where the pouch has not been completely sealed. .  Do not use Test Cards where the desiccant indicator has changed from blue to pink.  False-negative reactions may result due to deterioration of the improperly stored Test Cards.

- Do not use the Sample Diluent or Positive Control if they are discolored or turbid.  Discoloration or turbidity may be a sign of microbial contamination.  It is normal for the ImmunoCard Wash Buffer I to appear slightly turbid.

- The ImmunoCard Substrate I is light sensitive and should be stored in the dark.  The appearance of the reagent should be clear and colorless.  If the reagent exhibits any color or is turbid it should be discarded.

- Hold reagent vials vertically when dispensing drops to ensure consistent drop size and delivery.

STORAGE REQUREMENTS:

- The kit expiration date is indicated on the kit label.

- Store at 2 - 8 C and return the kit promptly to the refrigerator after each use.

At this facility, kits are stored:_______________________________________________

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

Each Test Card contains an internal control.  The appearance of a blue color in the CONTROL reaction port verifies the assay was active at the time of use, that sample was added and that there was adequate migration of the sample.

The reactivity of each new lot and each new shipment of the ImmunoCard Toxin A & B should be verified on receipt or before use.  Positive Control reagent is supplied with the kit.  The control is used to verify the reactivity of the other reagents associated with the assay, and is not intended to ensure precision at the analytical assay cut-off.  Sample Diluent is used for the Negative Control.  Additional tests can be performed with the controls to meet the requirements of local, state and/or federal regulations and/or accrediting organizations.  The results expected with the controls are described in the section on INTERPRETATION OF RESULTS.  The Test Card should not be used if control tests do not produce the correct results.  Proper results obtained with the Control Port, Positive Control and Negative Control (Sample Diluent) serve as indicators that the test was performed correctly, that the antibodies embedded in the membrane and the Enzyme Conjugate are active at the time of testing, and that the membrane supports proper sample flow.  Failure of the internal and external control to produce the expected results suggests the test was not performed correctly (ie, incorrect volume of reagents added; incorrect incubation temperature or times used or that reagents were not brought to room temperature prior to testing).  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented:________________________________________________________________

TEST PROCEDURE:

- Bring all test cards, reagents and samples to room temperature (20-26 C) before testing.  Remove the reagents from the kit box to warm.  Reagents may take up to 60 minutes to warm following refrigeration.  RETURN TO 2-8 C IMMEDIATELY AFTER USE.

- Label one test tube for each patient specimen to be tested.

- Add 200 (L of Specimen Diluent to a test tube using the dropper in the bottle.

- Add three drops of Enzyme Conjugate (150 (L) to each tube.

- Immediately add stool or control sample as follows:

- Liquid/Semi-solid stools – Mix the stool specimen thoroughly.  Using a clean transfer pipette (included with the kit), draw the stool to the first mark from the end of the tip (25 (L).  Dispense the stool into the Diluent/Conjugate mixture.  Use the same pipette to mix the stool by gently aspirating, then dispensing the suspension several times.  Vortex the final mixture for 10 seconds.  Leave the pipette in the tube for use later in the test.*

- Formed/Solid stools – Mix the stool specimen thoroughly.  Using a wooden applicator stick, transfer a small portion (approximately 2 mm diameter) of the stool into the Diluent/Conjugate mixture.  Emulsify the specimen using the applicator stick, and then thoroughly mix with a vortex for 10 seconds.  Add a pipette to the tube and leave for later use in testing.*

- Optional control:  Using the container droppers, add 1 drop of Positive Control or Negative Control (Specimen Diluent) (provided with the kit) to the Diluent/Conjugate mixture.  Vortex for 10 seconds.  Add a pipette to the tube and leave for later testing.

*NOTE: The transfer of too little sample, or failure to mix and suspend the sample in the Diluent/Conjugate mixture completely may result in a false-negative test result.  Addition of too much stool sample may cause invalid results due to restricted sample flow.

6.	Let each diluted sample stand for 5 minutes to 24 hours at 20-26 C.  Seal tube containing

diluted sample or control if incubating beyond 6 minutes.

7.	Use 1 ImmunoCard Test Card for each sample or control.  When ready to perform testing, remove the ImmunoCard Test Card from its foil pouch.  Discard the pouch and desiccant. Label the device with the name of the patient or the control.

8.	Vortex each sample or control for 10 seconds before use.

9.	Add 150 (L of each diluted sample or control to each of the 2 sample ports on 1 test card (150 (L represents the second mark on the transfer pipette).

10.	Incubate for 5 ± 1 minutes at 20-26 C.  NOTE: At the end of incubation both reaction ports must appear completely wet.  If either reaction port is not completely wet, discard the Test Card and repeat the procedure.  Proceed with testing if sample remains in the sample ports yet both reaction ports are completely wetted.

11.	Hold the ImmunoCard Wash Buffer I vial vertically and dispense exactly 3 drops to each of the reaction ports.

12.	When the ImmunoCard Wash Buffer I has been completely absorbed, hold the vial vertically and add exactly 3 drops of ImmunoCard Substrate I.

13.	Incubate the test card at 20-26 C for 5 minutes.

14.	Visually read the results of each card within 30 seconds of the end of incubation.

This facility’s procedure for sample labeling is:_____________________________________ ___________________________________________________________________________

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

POSITIVE                 NEGATIVE                                      INVALID

READ ONLY UPPER REACTION PORTS TO INTERPRET RESULTS

NOTE: Disregard color development in the sample ports.  The color of the sample ports is not evaluated during testing.

Positive test result: Blue color in the TEST (upper right) and CONTROL (upper left) reaction ports.  The intensity of the blue color of the TEST port may vary from the bottom to the top of port.  A positive test result indicates that toxin A and/or B is/are present in the sample.  PLEASE NOTE: The TEST port may appear less blue than the CONTROL port.

Negative test result: Blue color in the CONTROL reaction (upper left) port only.  The TEST reaction (upper right) port should be colorless.  (The wetting of the membrane may make the TEST port appear to be slightly grey.)  A negative test result indicates that C. difficile toxins are absent or below the limit of detection of the assay.  Occasionally the TEST PORT (upper right) may show evidence of a hint of blue color in the left or right side of the port, with the rest of the port colorless.  This should be considered a negative test result.

Invalid test results:

- No detectable blue color in the CONTROL reaction port.  Failure of the CONTROL invalidates any test result.  The invalid test may be due to the failure of a reagent or the Test Card to perform properly, failure to add sample, failure of the sample to migrate, or failure to dilute the sample correctly leading to over inoculation of the Test Card.  Samples with extremely high levels of toxins may also present with a positive TEST yet negative CONTROL port.

- A blue ring on the plastic frame surrounding the TEST (upper right) port during the test procedure.

When invalid results are obtained, redilute the sample and repeat the test on a new card.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Report test results as Toxin A and/or B is not present.

Positive test:  Report test result as Toxin A and/or B is present.

This facility’s procedure for patient result reporting is:_ _________________________________ ____________________________________________________________________________________________________________________________________________________________

EXPECTED VALUES:

The frequency of antibiotic-associated diarrhea caused by C. difficile is dependent on several factors, including patient population, type of institution and epidemiology.  The reported incidence of C. difficile-associated disease in patients suspected of having antibiotic-associated diseases is 15-20%, although different facilities may find positive rates above or below this range.  A positive result in an uninfected patient is not expected.

ASSAY SPECIFICITY

Stool samples inoculated with the following microbial agents (to a final sample concentration of ≥ 1 x 108 organisms/mL) do not react with ImmunoCard Toxins A & B: Adenovirus 40, Adenovirus 41, Aeromonas hydrophila, Bacillus cereus, Bacillus subtilis, Bacteroides fragilis, Campylobacter coli, Campylobacter jejuni, Candida albicans, Clostridium butyricum, nontoxigenic C. difficile (10 strains), Clostridium perfringens, Clostridium septicum, Clostridium sordellii strain 9714, Clostridium sporogenes, Enterobacter cloacae, Escherichia coli (four strains including two O157:H7 strains), Helicobacter pylori, Klebsiella pneumoniae, Peptostreptococcus anaerobius, Porphyromonas asaccharolytica, Proteus vulgaris, Pseudomonas aeruginosa, Rotavirus, Salmonella typhimurium, Serratia liquefaciens, Shigella dysenteriae, Shigella flexneri, Shigella sonnei, Staphylococcus aureus, Staphylococcus aureus Cowan I, Staphylococcus epidermidis, Streptococcus faecalis, Vibrio parahaemolyticus, Yersinia enterocolitica.  The only non-C. difficile microorganism reactive with ImmunoCard Toxins A & B was C. sordellii VPI 9048.  This strain produces toxin A and B homologs HT and LT.

TESTs for Interfering substances

The following substances had no effect on test results when present in stool in the concentrations indicated: fecal fat (4.8% w/v), whole blood (40% v/v), mucin (3.5% w/v), Metronidazole (0.25% w/v), Vancomycin (0.25% w/v), barium sulfate (5% w/v), Imodium AD® (5% w/v), Kaopectate® Caplets (5 mg/mL), Pepto Bismol® (5% v/v).

Additional information regarding the performance characteristics of this product can be obtained by contacting Meridian’s Technical Support Department at (800) 343-3858 for assistance (USA) or your local distributor.

LIMITATIONS OF THE PROCEDURE:

1.  The test is qualitative and no quantitative interpretation should be made with respect to the intensity of the color when reporting the result.

2.  A positive test confirms the presence of toxins A and/or B only.  A physician must use the test results in conjunction with other diagnostic procedures and the patient’s clinical condition to establish a diagnosis of C. difficile-associated disease.  The levels of toxin do not always correlate directly with the presence or severity of disease.

3.  Isolates of C. sordellii may react with this test due to the immunological identity of the C. sordellii toxins.

4.  The performance of samples from pediatric patients has not been evaluated.

5.  Two distinct groups have been identified that can harbor C. difficile asymptomatically at very high rates.  Colonization rates of up to 50% and higher have been reported in infants and rates of up to 32% in cystic fibrosis patients.

6.  While relatively stable at 2-8 C, C. difficile toxins -- particularly at low concentrations -- easily degrade at room temperatures. The rate at which the toxins degrade differs from patient sample to patient sample and therefore the rate cannot be predicted. For this reason, best practice requires that sample be refrigerated or frozen within two hours of collection and the samples tested within the timeframes recommended in this insert. Do not accept samples that have not been collected, handled or transported properly.

PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience ImmunoCard® Toxins A & B.

REFERENCES:

Refer to Directional Insert – Meridian Bioscience ImmunoCard® Toxins A & B.