[image: image2.png]‘\‘ ‘,x o o
%’ Meridian

Bioscience, Inc.






Page 1 of 8

Procedure and revision number: SN750220 CLSI 12/16



Meridian Catalog #750220
Procedure:  ImmunoCard STAT!® HpSA
	Institution: 

	Address:

	Department:


	Prepared By:
	Date Adopted
	Supercedes procedure #:

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


	Distributed To:
	# of Copies
	Distributed To:
	# of Copies

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created.  The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might effect this procedure, are identified.  Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:
The ImmunoCard STAT! HpSA is a rapid in vitro qualitative procedure for the detection of Helicobacter pylori antigens in human stool.  The stool antigen detection is intended to aid in the diagnosis of H. pylori infection and the demonstrate loss of H. pylori stool antigen following treatment.  Conventional medical practice recommends that testing by any method to confirm loss of antigen be done at least four weeks following completion of therapy
ImmunoCard STAT! HpSA is a rapid lateral flow immunoassay that utilizes a dissociated monoclonal anti-H. pylori antibody as the capture and detector antibodies.  A diluted patient stool sample is dispensed into the sample port of the test device and the appearance of a pink-red line in the reading window next to the letter T after five minutes of incubation at room temperature indicates a positive result. 

SPECIMEN:

Preferred sample types:
1. Solid or formed, semisolid and liquid stool samples are acceptable.
Undesirable specimens:
1. Stool in transport media
2. Stool on swabs
3. Stool mixed with preservatives
Interfering substances:
The following substances were found to have no effect on results when present in stool at the concentrations indicated.
Tums® Antacid (5 mg/mL), Tagamet® (5 mg/mL), Prilosec® (5 mg/mL), Mylanta® Antacid (1:20), Pepto-Bismol® (1:20), Barium sulfate (5%), Whole Blood (50%), Leukocytes (50%), Mucin (3.4%), Stearic acid/palmitic acid (fecal fat) (4%), Hemoglobin (tarry stool) (12.5%)

Collection:
The specimen should be transported in an airtight container and stored at 2-8 C until tested.  The specimen should be tested as soon as possible, but may be held up to 72 hours at 2-8 C prior to testing.  If testing cannot be performed within this time frame, specimens should be frozen immediately upon receipt and stored frozen (≤ -20 C) until tested.  Specimens may be frozen and thawed twice.  Mix stool samples thoroughly (regardless of consistency) before testing.
This facility’s procedure for specimen collection is: _________________________________ _________________________________________________________________________
This facility’s procedure for transporting specimens is: ______________________________

 _________________________________________________________________________
This facility’s procedure for rejected specimens is: _________________________________ _________________________________________________________________________
MATERIALS AND EQUIPMENT:

Materials:

1. ImmunoCard STAT! HpSA Test Devices, in individual foil pouches with a desiccant
2. Sample Diluent, in a plastic dropper vial.  Use as supplied.
3. Positive Control, in a plastic dropper vial.  Use as supplied.
4. 100 µL transfer pipettes
Materials Not Provided:
1. Disposable latex gloves that should be used during the handling of the fecal samples as they are considered potentially hazardous material
2. Vortex for suspending the stool specimen in the Sample Diluent (Optional)
3. Interval timer
Preparation:

1. Bring specimens and reagents to room temperature (20º-26º C) before testing. 

PERFORMANCE CONSIDERATIONS:
1. All reagents are for in vitro diagnostic use only.
2. Patient specimens may contain infectious agents and should be handled and disposed of as potentially biohazardous.
3. Do not interchange reagents from different kit lot numbers.  Do not use kit components beyond the labeled expiration date of the kit.
4. Allow kit components and specimens to reach the room temperature (20-26 C) before performing a test, as cold reagents and/or specimens may decrease assay sensitivity.  Reagents may take 20-30 minutes to warm following refrigeration.
5. Stool samples must be mixed thoroughly (regardless of consistency) to ensure a representative sample prior to sampling.
6. Inspect Test Devices before removing the foil pouch.  Do not use Test Devices that have holes in the foil pouch or where the pouch has not been completely sealed.  False negative reactions may result due to deterioration of the improperly stored Test Device.
7. Do not use the Sample Diluent or Positive Control if turbid.  Turbidity may be a sign of microbial contamination.
8. The Positive Control should be handled as potentially infectious even though it contains inactivated H. pylori.
9. Hold reagent vials vertically when dispensing drops to ensure consistent drop size and delivery.
10. Do not deviate from the method described here or falsely positive or falsely negative results may occur.
11. Test instructions should be thoroughly read before performing any testing.
12. Do not use a device if its pouch was punctured prior to use.
13. On occasion, particulate matter may initially interfere with sample flow.  In cases where the Test Device does not readily absorb the diluted specimen, gently touch the bottom of the sample port with an applicator stick, moving the stool solid particle that might prevent the absorption.  Alternatively, a new aliquot of the sample can be withdrawn from the Diluent and retested.
14. Caution: Federal law restricts this device to sale by or on the order of a physician.
STORAGE REQUIREMENTS:
Store kit refrigerated 2º to 8º C and return kit to the refrigerator promptly after each use. DO NOT FREEZE.
At this facility, kits are stored: _______________________________________________ 
QUALITY CONTROL:
This test should be performed per applicable local, state, or federal regulations or accrediting agencies.
The reactivity of ImmunoCard STAT! Test Devices should be verified upon receipt using the external Positive and Negative Control reagents provided in the kit.  The number of additional tests performed with the external controls will be determined by the requirements of local, state or national regulations or accrediting agencies.   

Internal controls:  Internal controls are contained within the Test Device and therefore are evaluated with each test. 

1. A colored band appearing at the control line serves as a positive control and indicates the test has been performed correctly, that sample was added, that it flowed properly, and that the test reagents were active at the time of use.  
2. A clear background around the control or test lines serves as a negative control.  A background that obscures the reading of results invalidates the test and is an indication of reagent deterioration, inappropriate sample or improper test performance.

External controls:  For laboratories performing CLIA-WAIVED tests, the reactivity of each new kit lot or shipment of ImmunoCard STAT! HpSA Test Devices must be verified upon opening each new kit using the external Positive and Negative Control reagents provided in the kit.  Each untrained operator must test a positive and negative control at least once with each 20-test kit.  

The results expected with the Controls are described in the section on INTERPRETATION OF RESULTS.  The Test Devices should not be used if control tests do not produce the correct results.  Failure to achieve the expected results indicates either that the Test Devices are defective or that the test was not performed correctly.  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.  The Positive and Negative Control reagents are manufactured in an aqueous solution matrix.  Although specimen matrix interference has not been observed with this assay, the aqueous matrix of the controls may not adequately control for specimen matrix effects.  The National Committee for Clinical Laboratory Standards (NCCLS) guidelines recommend that matrix control materials should be used when available.  If the user wishes to comply with the guideline, the user will need to provide control material in matrix.

QC Testing Frequency and Documentation:

For this facility, External QC is run: _____________________________________________ 

Results of External QC and action(s) taken when control results are unacceptable are documented: ________________________________________________________________

SPECIMEN PREPARATION:

1. Liquid or Semi-solid stools – Unscrew the red cap from the Sample Diluent vial (red capped vial).  Use a clean calibrated transfer pipette (supplied with the kit) to draw the mixed sample to the second mark from the pipette tip (100 µL).  Dispense the sample into the Sample Diluent vial.  Use the same transfer pipette to mix the diluted sample thoroughly, but gently, by squeezing the pipette bulb three times.  Recap the vial tightly and mix thoroughly but gently by swirling the contents for 15 seconds.  Alternatively, mix for 15 seconds using a vortex mixer.  NOTE: Care should be taken when pipetting semisolid stool.  The addition of less than 100 µL of stool may cause a false-negative test.  The addition of more than 100 µL of stool may cause invalid results due to restricted sample flow.
[image: image1.emf]
Figure of 100 µL pipette
2. Formed/Solid stools – Unscrew the red cap of the Sample Diluent vial (red capped vial).  Use the white plastic applicator stick in the red cap to collect a small portion of stool (5-6 mm pellet).  Transfer the pellet to the Sample Diluent vial.  Recap the vial tightly and mix thoroughly but gently by swirling the contents of the vial for 15 seconds.  Alternatively, mix for 15 seconds using a vortex mixer.  Wooden applicator sticks may also be used to transfer solid stool to the Sample Diluent.  NOTE: The transfer of too little stool, or failure to mix and suspend the stool in Sample Diluent completely may result in a false-negative test result.   Care should be taken to transfer no less and no more than the amount indicated.  The addition of more than 100 µL of stool may cause invalid results due to restricted sample flow.

PROCEDURE: (QUALIFIES AS CLIA WAIVED PROCEDURE)
A. Test
1. Bring all test devices, reagents and samples to room temperature (20-26 C) before testing.
2. Use 1 ImmunoCard STAT! Test Device for each patient sample.
3. Remove the ImmunoCard STAT! Test Device from its foil pouch.  The Test Device is marked to indicate where test and control lines will appear.  The round window marked with an arrow is the test window where sample is added. 
4. Label the device with the patient’s name.  Prepare the specimen according to the instructions in the SPECIMEN COLLECTION AND PREPARATION section above.
5. Hold the diluted specimen vial upright and tap the bottom gently on the countertop before proceeding.
6. Cover the top of the diluted sample vial with absorbent paper to avoid splatter.
7. Break off the red tip on the outside of the red cap.  (Do not break off the white applicator stick on the inside of the cap.)
8. Hold the vial upside down and dispense 4 drops of diluted sample into the round window (at arrow) of the Test Device.  Do not touch the tip of the vial to the Test Device.
9. Set a timer and incubate the test at 20-26 C for 5 minutes.
10. At the end of 5 minutes, read the results within 1 minute.  See the INTERPRETATION OF RESULTS section below for a description of positive and negative test results.
B. Controls
Positive and Negative Controls are designed to show all reagents are reactive, specific, and capable of producing the expected results.
1. Bring all control reagents to 20-26 C before testing.
2. Use 1 ImmunoCard STAT! Test Device each for a Positive and Negative Control.  Label each device with the control to be tested.
3. Hold reagent vials upside down to dispense reagents.  
4. Add 4 drops of the Positive Control to the test window (at arrow) of 1 device.  Do not allow the tip of the vial to touch the sample port.
5. Break off the red tip on the outside of the red cap of an unused vial of Sample Diluent.  
6. Dispense 4 drops of the Sample Diluent to the test window (at arrow) of another Test Device.
7. Set a timer and incubate the tests at 20-26 C for 5 minutes.
8. After 5 minutes, read the results within 1 minute of test completion.

This facility’s procedure for patient preparation is: _____________________________ ______________________________________________________________________
This facility’s procedure for sample labeling is: ________________________________ _______________________________________________________________________ 

CALIBRATIONS:

There are no calibrations associated with this procedure.
CALCULATIONS:

There are no calculations associated with this procedure.
INTERPRETATION OF RESULTS:
Negative test result: Only one BLUE colored band (Control Line) appears across the central window of the device close to the letter “C”.  (H. pylori antigens are absent or below the level of detection.)  No other bands should be seen.  The background should not interfere with reading the test.
Positive test result: In addition to the BLUE band (Control Line), a distinguishable PINK-RED band (Test Line) also appears across the central window of the device close to the letter “T”.  The intensity of the band will vary depending on the antigen concentration in the specimen.  Any pink-red line, even very weak, must be considered as a positive result.  (A positive test line indicates that H. pylori antigens are in the specimen.)   The background should not interfere with reading the test.

Invalid test results: 

1. The BLUE band (Control Line) is absent, with or without a visually detectable PINK-RED band (Test Line), 

2. A PINK-RED band appears at the letter “T” in the window after six minutes, or there is a line at this position of another color other than pink-red,
3. No Control Line band appears close to the letter “C”.  (The test is invalid since a shift in or absence of the control line indicates that the test procedure was performed improperly or that deterioration of the reagents has occurred.)  
If any test is difficult to interpret, the test should be repeated with the same sample to eliminate the potential for error.  Obtain a new sample and retest when the original sample repeatedly produces unreadable results.

In the event this test becomes inoperable, this facility’s course of action for patient samples is: _______________________________________________________________________

REPORTING OF RESULTS: 

Negative for H. pylori: Report test result as, “H. pylori antigen is absent or below the level of detection.”  
Positive for H. pylori: Report test result as, “H. pylori antigen detected.”
This facility’s procedure for patient result reporting is: __________________________________ ____________________________________________________________________________________________________________________________________________________________   
EXPECTED VALUES:
Studies on the epidemiology of H. pylori have shown that this organism is present worldwide. Gastritis caused by H. pylori has been shown to correlate with age, ethnic background, family size and socioeconomic class.  The prevalence of H. pylori infection in a given population can vary from 20% to 90%.  In patients diagnosed with duodenal ulcers, however, it has been shown in every age group to be approximately 80%. Currently recommended eradication treatments have an efficacy rate between 75% and 90%. 
The ImmunoCard STAT! HpSA test detects the presence of H. pylori antigens in human stool.  Expected values for a given population should be determined for each laboratory.  The rate of positives may vary depending on geographic location, method of specimen collection, handling and transportation, test employed and general health environment of patient population under study.
ASSAY SPECIFICITY:

The specificity of ImmunoCard STAT! HpSA was tested utilizing the following bacterial, viral and yeast strains.  Positive and negative stools were spiked with ≥ 1 x 107 bacteria or yeast or 1 x 105 viruses.  None of the microorganisms tested yielded a positive result in the negative stool or interfered with detection of the positive stool.  Both the negative and positive stool was positive when spiked with Helicobacter pylori strain 43504.

Adenovirus Type 2, Adenovirus Type 40, Coxsackie Type B1, Coxsackie Type B6, Echovirus Type 22, Feline calicivirus, Rotavirus, Bacillus species, Borrelia burgdorferi, Aeromonas hydrophila, Campylobacter coli, Campylobacter jejuni, Candida albicans, Citrobacter freundii, Clostridium perfringens, Clostridium difficile (2), Enterobacter cloacae, Enterococcus faecalis (2), E. coli (2), E. coli O157:H7 (2), E. fergusonii, Helicobacter felis, Hemophilus influenzae, Klebsiella pneumoniae, Proteus vulgaris, Pseudomonas aeruginosa, Salmonella dublin, Salmonella (Group B), Salmonella hilversum, Salmonella minnesota, Salmonella typhimurium, Staphylococcus aureus, Staphylococcus aureus (Cowan I), Staphylococcus epidermidis, Serratia liquefaciens, Shigella boydii, Shigella dysenteriae, Shigella flexneri, Shigella sonnei, Yersinia enterocolitica
LIMITATIONS OF THE PROCEDURE:
1. The test is qualitative and no quantitative interpretation should be made with respect to the intensity of the positive line when reporting the result.
2. Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.
3. Antimicrobials, proton pump inhibitors and bismuth preparations are known to suppress H. pylori, and ingestion of these prior to H. pylori testing (culture, histology, rapid urease, UBT, antigen) may cause false-negative results.  If a patient has ingested these compounds within two weeks prior to performing the ImmunoCard STAT! HpSA test, a false-negative result may occur.  In such cases, the test should be repeated on a new specimen obtained two weeks after discontinuing treatment.  A positive result for a patient ingesting these compounds within two weeks prior to performing the ImmunoCard STAT! HpSA test, should be considered accurate.
4. Failure to add sufficient stool to the Specimen Diluent may result in a falsely negative test result.  Addition of too much stool may result in invalid test results due to the inhibition of proper sample flow.
5. Overincubation of tests may lead to false-positive test results.  Incubating tests at reduced temperatures or times may lead to falsely negative results.
6. Performance characteristics have not been established for watery diarrheal stools.  Watery stools composed mainly of fluid with little or no solid matter may give false negative test results.
7. Laboratories performing CLIA WAIVED tests: If the user laboratory modifies the test system instructions, then the test is considered high complexity and subject to all applicable CLIA requirements.
PERFORMANCE CHARACTERISTICS:
Refer to Directional Insert – Meridian Bioscience ImmunoCard STAT!® HpSA
REFERENCES:
Refer to Directional Insert – Meridian Bioscience ImmunoCard STAT!® HpSA
Meridian Bioscience

3471 River Hills Drive

Cincinnati, OH 45244 USA



Ph:  (800) 343-3858, (513) 271-3700

SN750220 CLSI 12/16

[image: image2.png]