Procedure and revision number: SN11179 CLSI 01/20

Procedure:  TRU Legionella

Meridian Catalog #751930

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The TRU Legionella assay is an in vitro, rapid, lateral-flow immunoassay for the qualitative detection of Legionella pneumophila serogroup 1 antigen in human urine specimens.  It is designed to test specimens from patients with symptoms of pneumonia.  Test results are to be used as an aid in diagnosis of Legionella pneumophila serogroup 1 infection.  A negative result does not preclude infection with Legionella pneumophila serogroup 1.  Test results are to be used in conjunction with information obtained from the patient’s clinical evaluation and other diagnostic procedures.

TRU Legionella is a single use capture immunoassay to detect Legionella pneumophila serogroup 1 in human urine specimens.  The test consists of a Conjugate Tube, a Test Strip, and Sample Diluent/Negative Control.  The Conjugate Tube contains a lyophilized bead of colloidal gold-linked polyclonal antibody to Legionella pneumophila serogroup 1 detector antibody and polyclonal donkey anti-chicken IgY control detector antibody.  The Test Strip carries a nitrocellulose membrane with dried capture antibodies placed at a designated Test Line for Legionella.  The Test Strip holder caps the Conjugate Tube during testing and subsequent disposal to reduce exposure to potential pathogens.

The conjugate bead is first rehydrated in the Conjugate Tube with the Sample Diluent/Negative Control.  Patient sample is then added, the contents mixed and the Test Strip added. Legionella pneumophila serogroup 1 antigens, if present, bind to the antibody-colloidal gold conjugate.  When the sample migrates up the Test Strip to the Test Line, the antigen-conjugate complex is bound to the capture antibody, yielding a pink-red line.  When no antigen is present, no complexes are formed and no pink-red line appears at the Test Line.  An internal control line helps determine whether the test has been executed properly, if kit reagents are performing appropriately and that adequate flow has occurred through the Test Strip during a test run.  A visible pink-red line at the Control position of the Test Strip should be present each time a specimen or control is tested.  If no pink-red control line is seen, the test is considered invalid.

SPECIMEN:

Preferred sample types:

Human urine samples, unpreserved or preserved with boric acid

Undesirable specimens:

Boiled urine

COLLECTION:

Human urine samples, unpreserved: Samples should be received in an airtight transport container and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to seven days at 2-8 C.  Samples that will not be tested within seven days should be frozen immediately upon receipt and stored at ≤ -20 C until tested.  Samples may be frozen and thawed up to two times after storage at ≤ -20 C prior to testing.

Human urine samples, boric acid preserved: Samples should be received in an airtight Boric Acid Urine Tube and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to seven days at 2-8 C.  (See SPECIMEN PREPARATION section for instructions on boric acid preserved urine.)  Samples that will not be tested within seven days should be frozen immediately upon receipt and stored at ≤ -20 C until tested.  Samples may be frozen and thawed up to two times after storage at ≤ -20 C prior to testing.

This facility’s procedure for specimen collection is: _____________________________________ ______________________________________________________________________________

This facility’s procedure for transporting specimens is: __________________________________

______________________________________________________________________________

This facility’s procedure for rejected specimens is: _____________________________________ ______________________________________________________________________________

MATERIALS AND EQUIPMENT:

MATERIALS PROVIDED:

The maximum number of tests obtained from this test kit is listed on the outer box.

- Test Strip: A test strip attached to a plastic holder enclosed in a foil pouch with desiccant.  The test strip carries anti-Legionella pneumophila serogroup 1 capture antibody for the test line.  The holder is used to stopper the Conjugate Tube.  The strip frame portion of the holder indicates where test and control lines should appear.  Store the pouch at 2-27 C when not in use.  Do not use a device if the desiccant indicator (line in center of desiccant) has changed from blue to pink.

- Conjugate Tube: A capped plastic tube containing a conjugate bead.  The tube is enclosed in a foil pouch.  The conjugate consists of gold-conjugated anti-Legionella pneumophila serogroup 1, which serves as the detector antibody and polyclonal donkey anti-chicken IgY control detector antibody.  Store the foil pouch at 2-27 C when not in use.  Do not store in the freezer.  Do not remove the cap until just before use.

- Sample Diluent/Negative Control: A buffered protein solution provided in a plastic vial.  Sodium azide (0.094%) added as a preservative.  Use as supplied.  Store at 2-27 C when not in use.

- Positive Control: Inactivated Legionella pneumophila serogroup 1 in a buffered diluent containing sodium azide (0.094%) as a preservative.  The reagent is supplied ready to use.  Store at 2-27 C when not in use.

- Plastic transfer pipettes (with 100, 200 and 300 µL volume marks).

MATERIALS NOT PROVIDED:

1.	Disposable latex gloves (urine samples are considered potentially biologically hazardous material.)

2.	Vortex for suspending the specimen in the Sample Diluent (optional)

3.	Interval timer

4.	Marking pen

5.	Optional: Urine preservative containing boric acid, such as the BD Vacutainer® C&S Preservative Plus Urine Tube

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- DO NOT interchange the Test Strips, Conjugate Tubes and Positive Control reagents from different kit lot numbers.  The Sample Diluent/Negative Control can be interchanged between kits providing the reagents are within their expiration periods.  Do not use any reagent beyond its labeled expiration date.

- Do not use components that lack a label, lot number, or an expiration date.

- Do not use any reagent if it is discolored or turbid.  Discoloration or turbidity may be a sign of microbial contamination.

- Allow reagents to warm to 19-27 C before use.

- All reagents should be mixed gently before use.

- The Positive Control reagent vial should be held vertically when dispensing drops to ensure consistent drop size and delivery.

- The transfer pipettes provided with the kit must be used for specimen preparation and transfer.  Use one per specimen.

- Any deviation below or above set incubation times may affect sensitivity and specificity and should be avoided.

- Test Strips and Conjugate Tubes are packaged in foil pouches that exclude moisture during storage.  Inspect each foil pouch before use.  Do not use Test Strips or Conjugate Tubes from pouches that have holes, where the pouch has not been completely sealed or where the desiccant indicator has changed from blue to pink.  The change in the desiccant color is an indicator the Test Device has been exposed to moisture.  False-negative reactions may result if Test Strips or Conjugate Tubes are exposed to moisture.

- All urine samples must be mixed thoroughly before testing to ensure a representative sample prior to pipetting.

- Legionella pneumophila serogroup 1 antigens may be relatively unstable.  Care should be taken to store samples as indicated in this procedure.  Even when samples are stored in the frozen state, the rate at which antigen deterioration occurs varies from sample to sample and cannot be predicted.

- Sample Diluent/Negative Control must be added to the Conjugate Tube within one minute after removing the cap from the tube.

SHELF LIFE AND STORAGE:

- The kit expiration date is indicated on the kit label.

- Store at 2-27 C when not in use.

At this facility, kits are stored: ______________________________________________________

CALIBRATION:

There are no calibrations associated with this procedure.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

At the time of each use, kit components should be visually examined for obvious signs of microbial contamination, freezing or leakage.  Do not use contaminated or suspect reagents.

Internal procedural controls: Internal procedural controls are contained within the Test Strip and therefore are evaluated with each test.

- A PINK-RED band appearing at the Control Line serves as a procedural control and indicates the test has been performed correctly, that proper flow occurred and that the test reagents were active at the time of use.

- A clean background around the Control or Test Lines also serves as a procedural control.  Control or Test Lines that are obscured by heavy background color may invalidate the test and may be an indication of reagent deterioration, use of an inappropriate sample or improper test performance.

External Control reagents should be tested according to the requirements of the laboratory or applicable local, state or accrediting agencies.

      - See section EXTERNAL CONTROL TESTS for instructions on performing these control tests.

- The reactivity of each new lot and each new shipment of TRU Legionella should be verified on receipt using external Positive and Negative Control reagents.  The number of additional tests performed with external controls will be determined by the requirements of local, state or federal regulations or accrediting agencies.

- The external controls are used to monitor reagent reactivity.  Failure of the controls to produce the expected results can mean that one of the reagents or components is no longer reactive at the time of use, the test was not performed correctly, or that reagents or samples were not added.  If the Positive or Negative External Controls fail, do not report test results to the clinician.

- The results expected with the Controls are described in the section on INTERPRETATION OF RESULTS.

- The kit should not be used if control tests do not produce the correct results.  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure.  If the control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

- Positive and Negative Control reagents manufactured for this assay are prepared in the matrix of the Sample Diluent/Negative Control, which may not mimic test specimens.  If control materials that are identical in composition to test specimens are preferred, the user can prepare those by diluting known positive and negative specimens in Sample Diluent/Negative Control according to the SPECIMEN PREPARATION section of this insert.

QC Testing Frequency and Documentation:

For this facility, External QC is run: __________________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented: ___________________________________________________________________

PROCEDURAL NOTES:

The TRU LEGIONELLA transfer pipette is diagrammed below.

The TRU LEGIONELLA Test Strip and Conjugate Tube are diagrammed below.

SPECIMEN PREPARATION:

Mix urine thoroughly prior to pipetting.  Bring specimens and reagents to room temperature (19-27 C) before testing.

1.	Human urine samples, unpreserved and preserved:

- Remove one Conjugate Tube from its foil pouch and discard the pouch.  Label the tube with the patient’s name.

- Remove the cap from the Conjugate Tube and discard the cap.

- Using a transfer pipette supplied with the kit, immediately add 100 µL (first mark from the tip of the pipette) of Sample Diluent/Negative Control to the Conjugate Tube.  Dispense directly into the center of the tube.  Vortex or swirl the contents of the   Conjugate Tube for 10 seconds. *

- Mix patient sample thoroughly.  Use one of the transfer pipettes supplied with the kit to mix the sample gently but thoroughly by squeezing the pipette bulb three times.  Alternatively, mix for at least 10 seconds using a vortex mixer.

- Using the same pipette, draw 100 µL of specimen (first mark from the end of the pipette) and add it to the Conjugate Tube.

- Using the same pipette, mix the sample and conjugate thoroughly but gently by squeezing the pipette bulb three times.  Alternatively, mix for at least 10 seconds using a vortex mixer.  Discard the pipette.

*Warning: Dilution errors may affect test performance.  Failure to add sufficient urine sample to the Sample Diluent/Negative Control may result in falsely negative tests.  Failure to add the full amount of Sample Diluent/Negative Control may result in falsely positive tests.  Addition of too much sample may result in invalid test results due to the inhibition of proper sample flow.

This facility’s procedure for patient preparation is: __________________________________ ___________________________________________________________________________

This facility’s procedure for sample labeling is: _____________________________________ ___________________________________________________________________________

TEST PROCEDURE:

1.	Remove the Test Strip from its foil pouch and discard the pouch.  Visually inspect Test        Strip.  Do not use Test Strip if a green line is present.

2.	Insert the narrow end of the Test Strip into the Conjugate Tube and firmly press down on the cap to close the tube.

3.	Incubate at 19-27 C in an upright position in a rack for 20 minutes.

4.	Read the results on the test strip within 1 minute.  Do not read results beyond this period.  (NOTE: Remove the Test Strip from the Conjugate Tube if Test or Control Lines are difficult to read.  Recap the Conjugate Tube with the Test Strip holder and discard when testing is completed.)

EXTERNAL CONTROL TESTS:

1.	Bring all test components, reagents and samples to room temperature (19-27 C) before testing.

2.	Use 1 Conjugate Tube and 1 Test Strip for Positive Control testing and 1 Conjugate Tube and 1 Test Strip for Sample Diluent/Negative Control testing.

3.	Remove the Conjugate Tubes from their foil pouches and label accordingly.  Discard the pouches.

4.	Remove the caps from the Conjugate Tubes and discard the caps.

5.	Add exactly 5 drops of the Positive Control reagent to the Conjugate Tube marked for the Positive Control.  The drops should be dispensed directly into the center of the tube.

6.	Using 1 of the transfer pipettes supplied with the kit, add 200 µL (second mark from the end of the pipette tip) of Sample Diluent/Negative Control to the Conjugate Tube marked for the Negative Control.  Add directly to the center of the tube.

7.	Vortex or swirl the contents of the tubes for 10 seconds.

8.	Remove 2 Test Strips from their foil pouches and discard the pouches.

9.	Insert the narrow end of a Test Strip to each Conjugate Tube and firmly press down on the caps to close each tube.

10.	Incubate both tubes at 19-27 C for 20 minutes.

11.	Read the results on the test strip within 1 minute.  Do not read results beyond this period.  (NOTE:  Remove the Test Strip from the Conjugate Tube if test or control lines are difficult to read.  Recap the Conjugate Tube with the Test Strip holder and discard when testing is completed.)

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Negative test: A PINK-RED band at the Control Line position.  No other bands are present.

Positive test: PINK-RED band of any color intensity at the Control and Legionella Test Line positions.  The color of the Test Line can be lighter than that of the Control Line.  Test Lines may appear strongly visible or may appear less strongly visible.

Invalid test results:

1.	No band at the designated position for the Control Line.  The test is invalid since the absence of a control band indicates the test procedure was performed improperly or that deterioration of reagents has occurred.

2.	A PINK-RED band appearing at the Test Line position of the device after 21 minutes of incubation, or a band of any color other than PINK-RED.  Falsely positive results may occur if tests are incubated too long.  Bands with colors other than PINK-RED may indicate reagent deterioration.

If any result is difficult to interpret, the test should be repeated with the same sample to eliminate the potential for error.  Obtain a new sample and retest when the original sample repeatedly produces unreadable results.

In the event that this test becomes inoperable, this facility’s course of action for patient samples is: ____________________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Report test results as “No L. pneumophila serogroup 1 antigens detected.  A negative result does not exclude infection with Legionella pneumophila serogroup 1 nor does it rule out other microbial-caused respiratory infections or disease caused by other serogroups of Legionella pneumophila.”

Positive test:  Report test result as “L. pneumophila serogroup 1 antigens detected.  This result does not rule out coinfection with other pathogens.”

This facility’s procedure for patient result reporting is: __________________________________ ____________________________________________________________________________________________________________________________________________________________

EXPECTED VALUES:

Legionella pneumophila serogroup 1 is an important cause of travel, community and hospital acquired pneumonia worldwide.  Legionnaires’ disease (LD) is known to occur sporadically and in outbreak settings.  Recent data indicates that 0.5-5.0% of adults hospitalized for pneumonia has LD.  Approximately 8,000-18,000 cases of LD occur each year in the USA.2, 3

ASSAY SPECIFICITY

The following L. pneumophila stock cultures from different sources were tested and produced positive reactions at a minimum concentration of 4.8 x 106 CFU/mL with the TRU LEGIONELLA assay:

Pontiac subgroup: CCUG 13395; NCTC 12024, Allentown 1; NCTC 12006, Benidorm; CCUG 33058, Knoxville; NCTC 12007, France.

Non-Pontiac subgroup: NCTC 12008, OLDA; NCTC 12098, Camperdown; NCTC 12025, Heysham; NCTC 12009, Oxford.

Crossreactivity studies were performed with positive and negative urine specimens inoculated with bacterial or fungal organisms to a final concentration of 1.2 x 108 CFU/mL or viral concentration greater than 1 x 105 TCID50/mL.  None of the following organisms in urine reacted with TRU LEGIONELLA:

Alcaligenes faecalis, Bacillus cereus, Bacillus subtilis, Candida albicans, Candida glabrata, Candida parapsilosis, Citrobacter freundii, Enterobacter aerogenes, Enterobacter cloacae, Enterococcus faecalis, Enterococcus faecium, Escherichia coli, Haemophilus influenzae, Klebsiella pneumoniae, Legionella bozemanii, Legionella dumoffii, Legionella feeleii, Legionella gormanii, Legionella longbeachae, Legionella micdadei, Legionella pneumophila serogroup 2 Togus, Legionella pneumophila serogroup 3 Bloomington, Legionella pneumophila serogroup 4, Legionella pneumophila serogroup 5, Legionella pneumophila serogroup 6, Morganella morganii, Moraxella osloensis, Mycoplasma pneumoniae, Nocardia asteroides, Proteus mirabilis, Proteus vulgaris, Pseudomonas aeruginosa, Serratia liquefaciens, Serratia marcescens, Staphylococcus aureus, Staphylococcus epidermidis, Staphylococcus saprophyticus, Streptococcus groups A, B, D, F, G, Streptococcus pneumoniae, Adenovirus, Coxsackievirus, Influenza A, Influenza B, Parainfluenza Virus, Respiratory Syncytial Virus A, Respiratory Syncytial Virus B.

## Tests for Interfering Substances:

The following substances, at the specified saturated solvent/diluent concentrations, do not interfere with the test results in the final concentrations listed:

Amphotericin B (0.22 mg/mL), Antihistamine (0.22 mg/mL), Ascorbic acid (1.0 mg/mL), Beet root (0.01%), Bilirubin (0.2 mg/mL), Boric acid (2.63 mg/mL), Caffeine, purified (0.4%), Chlorophyll (0.81 mg/mL), Ciprofloxacin (0.22 mg/mL), Cold and flu tablets (50 mg/mL), Cough drops (0.22 mg/mL), Cough syrup (0.20 mg/mL), Decongestant (0.22 mg/mL), Erythromycin (0.067 mg/mL), Glucose (20 mg/mL), Itraconazole (0.22 mg/mL), Miconazole (5%), Oxalic acid (0.01%), Prednisone (0.22 mg/mL), Protein (BSA) (5 mg/mL), Rifampicin (0.09 mg/mL), Tobacco, purified (0.4%), Urea (20 mg/mL), Vaginal contraceptive gel with nonoxynol-9, 4% (5%), Water-based personal lubricant (KY Jelly) (5%), White blood cells (10%), Whole blood (10%).

LIMITATIONS OF THE PROCEDURE:

      - The test is qualitative and no quantitative interpretation should be made with respect to the intensity of the positive Test Line when reporting the result.

      - TRU Legionella cannot be used as the sole means of determining Legionella pneumophila infection.  Test results must be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures.

      - Overincubation of tests may lead to an increase in false-positive test results.  Conversely, incubation for periods less than those defined in this insert can result in an increase in false-negative tests.  Follow incubation times defined in this insert.

      - TRU Legionella detects both viable and non-viable Legionella pneumophila serogroup 1.  The appearance of TRU Legionella positive tests depends on the antigen load in the specimen; therefore a TRU Legionella positive test may not always correlate with the results of bacterial culture.

      - The antibodies used in the test may not detect all antigenic variants or new strains of Legionella pneumophila serogroup 1.

      - A negative test result does not exclude infection with Legionella pneumophila serogroup 1 nor does it rule out other microbial-caused respiratory infections or disease caused by serogroups of Legionella pneumophila other than serogroup 1.  A positive test result does not rule out coinfection with other microbes.

      - The TRU Legionella assay is prescription use only.

      - The TRU Legionella is intended for use with urine specimens.  Use of this assay is not recommended with other specimens.

      - A positive results implies the presence of antigen to Legionella pneumophila serotype 1, however, all test results should be reviewed in light of other clinical data by the physician.

      - The TRU Legionella assay is not intended for monitoring therapy.

      - Performance characteristics are not established for patients younger than 21 years of age.

      - The performance of the TRU Legionella assay has only been established for manual reading and/or visual result determination.

## PERFORMANCE CHARACTERISTICS:

Refer to Package Insert – Meridian Bioscience TRU Legionella®

## REFERENCES:

Refer to Package Insert – Meridian Bioscience TRU Legionella®