Procedure and revision number: 755725  Rev. 3854-6   CLSI 11/18

Procedure:  ImmunoCard STAT! Mono

Meridian Catalog 755725

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might effect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The ImmunoCard STAT! Mono Test is intended for the qualitative detection of infectious mononucleosis heterophile antibodies in serum, plasma or whole blood as an aid in the diagnosis of infectious mononucleosis.

The ImmunoCard STAT! Mono Test uses color immunochromatographic dipstick technology with bovine erythrocyte extract coated on the membrane.  In the test procedure, serum, plasma or whole blood is mixed with the Diluent.  Then the Test Stick is placed in the mixture and the mixture migrates along the membrane.  If the specific IM heterophile antibody is present in the sample, it will form a complex with the bovine erythrocyte extract conjugated color particles.  The complex will then be bound by bovine erythrocyte extract immobilized on the membrane and a visible blue Test Line will appear to indicate a positive result.

SPECIMEN:

Preferred sample types:

###### Serum, Plasma or Whole Blood Sample collected in a tube containing EDTA or heparin as an anticoagulant.

###### Fingertip Whole Blood

Undesirable specimens:

###### Serum, Plasma or Whole Blood Sample collected in a tube containing an anticoagulant other than EDTA or heparin.

###### Specimen Collection and Preparation:

###### Serum, Plasma or Whole Blood Sample

Obtain specimens by acceptable medical technique.  Collect whole blood samples using a tube containing EDTA or heparin as an anticoagulant.  Other anticoagulants have not been tested.  Serum and plasma specimens may be refrigerated (2 - 8C) and tested within 48 hours; serum and plasma specimens held for longer times should be frozen (below -10C) and tested within 3 months.  Test whole blood specimens within 24 hours.  Specimens must be at room temperature (15 – 30C) when tested.

###### Fingertip Whole Blood

Hold the capillary pipette horizontally , and touch the tip of the pipette to the drop of blood on the patient’s finger until it fills completely to the line.  Note:  Filling is automatic; never squeeze the pipette bulb while collecting sample.

This facility’s procedure for patient preparation is: ________________________________.

This facility’s procedure for sample labeling is: ___________________________________.

This facility’s procedure for specimen collection is: _________________________________ __________________________________________________________________________

This facility’s procedure for transporting specimens is: ______________________________

__________________________________________________________________________

This facility’s procedure for rejected specimens is: _________________________________ __________________________________________________________________________

MATERIALS AND EQUIPMENT:

Materials:

- 25  Test Sticks in a container

- 25  Test Tubes

- 25  Transfer pipettes

- 27 Capillary Pipettes

- 1 Diluent (contains buffer with 0.2% sodium azide)

- 1 Mono Positive Control (contains rabbit anti-beef stroma in tris buffer with 0.2% sodium azide and 0.05% gentamycin sulfate preservatives)

- 1 Mono Negative Control (contains goat albumin in tris buffer with 0.2% sodium azide)

- 1 Work Station

- 1 Directional insert

Note: Extra components (tubes, pipettes, capillary pipettes) have been provided for	your convenience.

Materials Required But Not Provided:

- A timer or watch

- Specimen collection containers

- Disposable latex gloves

PERFORMANCE CONSIDERATIONS:

- Caution:  Federal Law restricts sales of this device to or on the order of a licensed practitioner.

- All reagents are for in-vitro diagnostic use only.

- Follow your laboratory safety guidelines in the collection, handling, storage and disposal of patient specimens and all items exposed to patient specimens.

- The Diluent and Controls contain sodium azide, which may react with lead or copper plumbing to form potentially explosive metal azide.  Large quantities of water must be used to flush discarded Diluent or Controls down a sink.

- Do not interchange or mix components from different kit lots.

STORAGE REQUIRMENTS:

- Store the Test Sticks and reagents tightly capped at 15 – 30C.

- Do not use the Test Sticks or reagents after their expiration dates.

At this facility, kits are stored: _______________________________________________

CALIBRATION:

Not applicable to this assay.

QUALITY CONTROL:

This test should be performed per applicable local, state, or federal regulations or accrediting

agencies.

###### External Quality Control

For external QC testing, use the controls provided in the kit.  Add one free falling drop of 	control to the Test Tube and then proceed in the same manner as with a patient sample. 	Quality Control requirements should be established in accordance with local, state and 	federal regulations or accreditation requirements. Minimally, Meridian Bioscience 	recommends that positive and negative external controls be run with each new lot and 	with each new untrained operator.  Some commercial controls may contain interfering 	additives.  The use of these controls is not recommended.

###### Internal Quality Controls

The ImmunoCard STAT! Mono Test provides two levels of internal procedural controls 	with each test procedure.

- The red Control Line is an internal positive procedural control.  The Test Stick must absorb the proper amount of test material and be working properly for the red Control Line to appear.

- A clear background is an internal negative procedural control.  If the test has been performed correctly and the Test Stick is working properly, the background will clear to give a discernible result.

If the red Control Line does not appear, the test may be invalid.  If the background does not clear and interferes with the test result, the test may be invalid.  If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated please contact Meridian's Technical Services Department at 1-800-343-3858 (US) or your local distributor.

QC Testing Frequency and Documentation

For this facility, external QC is run: ____________________________________________.

________________________________________________________________________.

Results of External QC and action(s) taken when control results are unacceptable are documented: _____________________________________________________________.

________________________________________________________________________.

TEST PROCEDURE:

- Addition of specimen

  - For serum, plasma or whole blood samples in tubes:  Use the Transfer Pipette provided and add one drop to the Test Tube.

  - For Fingertip blood:  To dispense all the patient sample, place the pipette into the test tube and squeeze the bulb.  Remove pipette and discard in the appropriate biohazard container.  NOTE:  If pipette does not fully dispense the patient’s sample, recollect the sample using a new capillary pipette.

- Slowly add 1 drop of Diluent to the bottom of the Test Tube.

- Mix

- Remove the Test Stick(s) from the container.  Re-cap the container immediately.

- Place the Absorbent End of the Test Stick into the treated sample.  Leave the Test Stick in the Test Tube.

- Read results at 5 minutes.  Positive results may be read as soon as the red Control Line appears.

Discard used test tubes and Test Sticks in the suitable biohazardous waste container.

For this facility, sample swabs, used test tubes and Test Sticks are disposed: __________________________________________________________________________.

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

###### Note:  A blue or red line, which appears uneven in color density, is considered a valid result.

###### Positive

A blue Test Line and a red Control Line is a positive result for the detection of infectious mononucleosis heterophile antibody.  Note that the blue line can be any shade of color and can be lighter or darker than the line in the picture.

###### Negative

A red Control Line but no blue Test Line is a negative result.  No infectious mononucleosis heterophile antibody has been detected.

###### Invalid

If after 5 minutes, no red Control Line appears or background color makes reading the red Control Line impossible, the result is invalid.  If this occurs, repeat the test on a new Test Stick or call Meridian Bioscience Technical Services.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:

____________________________________________________________________________

REPORTING OF RESULTS

Negative test for Mononucleosis Heterophile Antibody:  Report test result as “Infectious mononucleosis heterophile antibody not detected, or below the sensitivity of this test kit.”

Positive test for Mononucleosis Heterophile Antibody:  Report test result as “Infectious mononucleosis heterophile antibody detected.”

This facility’s procedure for patient result reporting is: ______________________________                                                                                                                                    __________________________________________________________________________

EXPECTED VALUES:

A heterophile antibody response is observed in approximately 80-90% of adults and children with EBV-caused IM.  This percentage drops to approximately 50% for children under four years of age.  While the incidence of IM reflects wide seasonal, ethnic and geographical variation, a large epidemiological study noted that the highest incidence of symptomatic IM occurred during late adolescence (15-24 years of age).

ASSAY REACTIVITY:

When compared to a commercially available latex particle agglutination test for infectious mononucleosis heterophile antibodies, the ImmunoCard STAT! Mono Test showed a sensitivity of 100% and a specificity of 90.3%. The overall agreement was 94.1%.

LIMITATIONS OF THE PROCEDURE:

- As with all diagnostic assays, the results obtained by this test yield data that must be used as an adjunct to other information available to the physician.

- The ImmunoCard STAT! Mono Test is a qualitative test for the detection of IM heterophile antibody.

- A negative result may be obtained from patients at the onset of the disease due to heterophile antibody levels below the sensitivity of this test kit.  If symptoms persist or intensify, the test should be repeated.

- Some segments of the population with acute IM are heterophile antibody negative.

## PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience ImmunoCard Stat! Mono

## REFERENCES:

Refer to Directional Insert – Meridian Bioscience ImmunoCard Stat! Mono