[image: image4.png]


Page 7 of 7

Procedure:  Curian HpSA
Procedure and revision number: SN11285  CLSI REV 07/20

Meridian Catalog: 760130
	Institution: 

	Address:

	Department:


	Prepared By:
	Date Adopted
	Supersedes procedure #:

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


	Distributed To:
	# of Copies
	Distributed To:
	# of Copies

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.
PRINCIPLE:

Curian HpSA, for use with the Curian Analyzer is a rapid, qualitative, fluorescent immunoassay for the detection of Helicobacter pylori antigen in human stool.  Test results are intended to aid in the diagnosis of H. pylori infection and to demonstrate loss of H. pylori antigen following treatment.  Accepted medical practice recommends that testing by any current method, to confirm eradication, be done at least four weeks following completion of therapy. Test results should be taken into consideration by the physician in conjunction with the patient history and symptoms.
The Curian HpSA assay consists of a test strip enclosed in a plastic frame (test card), positive control reagent, and Aioprep™ sample preparation device.  Curian HpSA is a lateral flow-based immunoassay for the direct detection of H. pylori antigens in human stool.  Curian HpSA utilizes monoclonal anti-H. pylori antibodies as the capture and detector antibodies.  A sample of the patient’s stool specimen is collected using the sample collection brush that is included in the Curian HpSA assay kit.  The brush is pushed through a metering insert, which is integral to the Aioprep device, to remove excess stool.  The brush is then pushed directly into the prefilled sample diluent.  The diluted sample is mixed and dispensed drop-wise into the sample port of the Curian HpSA test card.  The H. pylori antigen (if present) binds to the anti-H. pylori antibody conjugated to fluorescent particles, forming a complex.  As the sample moves through the test strip, the anti-H. pylori capture antibody, bound to the assay membrane at the test position of the strip, binds the complex and yields a test line.  When antigen is not present, a complex is not formed and a test line will not occur.  As the conjugate continues to move further up the test strip, a detector antibody forms another complex in the control area of the test strip.  A line at the control position of the test strip should be present each time a sample or external control is tested.  If the control line is not generated, adequate sample flow has not occurred, and the Curian Analyzer will consider the test invalid.

MATERIALS AND EQUIPMENT:
Materials Provided:
The maximum number of tests obtained from this kit is listed on the outer box.
1. Curian HpSA Test Card: A test strip enclosed in a plastic frame which is in a foil pouch with a desiccant.  The test card is supplied ready to use.
2. Curian HpSA Aioprep Sample Preparation Device/ Negative Control: A buffered protein solution containing blue dye and 0.1% ProClin.  The Aioprep device is fitted with a metering insert and dropper tip.  Supplied ready to use.
3. Curian HpSA Positive Control: Inactivated H. pylori purified antigen in a phosphate buffered solution containing 0.094% sodium azide.  Supplied ready to use. 

4. Curian HpSA Sample Collection Brushes
5. Wooden Applicators 
Materials Not Provided:
1.  Disposable gloves, powder free.
Equipment Not Provided:
1. Vortex mixer
2. Interval timer (Optional)
3. Curian Analyzer System, Meridian Bioscience, Inc. Catalog 610190
PERFORMANCE CONSIDERATIONS:
1. All reagents are for in vitro diagnostic use only.  
2. Store the kit at the temperature indicated on labeling when not in use.
3. Do not interchange sample collection brushes between assays.  Specimens must be sampled and prepared with the Curian HpSA sample collection brush and used with the Curian HpSA Aioprep device.  These components are not interchangeable with other assays.  

4. Handle and dispose of all human specimens as if they are biologically hazardous.
5. Inspect foil pouch before removing the test card.  Do not use test cards that have holes in the foil pouch or where the pouch has not been completely sealed.  
6. Do not use test cards where the desiccant indicator has changed from blue to pink. 
7. Inspect Curian HpSA Aioprep / Negative Control prior to use.  Tap the Aioprep to ensure the liquid is in the main chamber of the Aioprep.
8. Do not interchange positive control or test cards between kit lots. 
9. The Curian HpSA test card must be incubated outside of a laminar flow hood, this can be done either inside the Analyzer or on the lab bench.  
10. Do not mark over or near the barcode on the test card.
STORAGE REQUIRMENTS:

1. The expiration date is indicated on the kit label.  
2. Store the kit at 2-8 C as indicated on the labeling.  
3. Return all kit components to the indicated storage temperature after use. 

At this facility, kits are stored: _______________________________________________
CALIBRATION:

Not applicable to this assay.
OPERATION OF THE CURIAN ANALYZER 

The Curian Analyzer is an easy to use, menu driven analyzer system.  Instructions to complete testing are provided on the analyzer touchscreen and in the Curian Analyzer operator’s manual. 

1. Power-on the Curian Analyzer by pressing the power button located on the left side of the analyzer.  The Curian Analyzer will initialize and perform self-checks. 

2. Instrument Checks, Quality Control testing, and patient specimen testing are completed by navigating the Curian Analyzer menus.  User ID and Sample ID entries are required for Quality Control and specimen testing. 

3. Complete use and operation of the Curian Analyzer System by following the instructions in the Curian operator’s manual and on-screen prompts.  Refer to operator’s manual for detailed instructions for the Curian Analyzer.
QUALITY CONTROL:
There are four controls for the Curian HpSA test system.  Two internal controls, one for the assay and one for the analyzer, and two external controls, one for the assay and one for the analyzer. 

1. Analyzer Internal Control: Self Test check (automatic)

2. Analyzer External Control: Instrument Check (IC Check)

3. Assay External Controls: Positive and Negative Controls (QC test)

4. Assay Internal Control: Control line

INSTRUMENT CHECK TEST PROCEDURE (Analyzer External Control Test):
Note:  The Analyzer’s External Control (IC Check) test has a default schedule of 30 days.  
       The IC Check test will be required to be performed prior to running a QC test or a patient test if this has not been performed within this timeframe.  The user will be locked-out of running a patient test if this test is overdue. 

1. Power-on the Curian by pressing the power button located on the side of the analyzer.  The Curian will initialize and perform Self Test checks. 

2. From the HOME screen, select “TEST”, and then select “Instrument Check”. 

3. Insert the Fluorescent IC Card into the drawer.  A pass or fail result will be displayed once the analysis is complete.
If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure.  Refer to operator’s manual for detailed instructions for the Analyzer. 
QUALITY CONTROL TEST PROCEDURE (Assay External Quality Control Test)

       Note:  Good Laboratory Practice guidelines recommend the use of control material.  Users should follow the appropriate federal, state, and local guidelines concerning the running of external quality controls. 

1. Bring all test components and reagents to 19-27 C before testing.  

2. Use 1 Curian test card for each QC sample control.  When ready to perform testing, remove the test card from its foil pouch.  
3. Discard the pouch and desiccant.  Do not use test cards where the desiccant indicator has changed from blue to pink. 
4. Label the test card with control identification information.  Do not mark over or near the barcode on the test card
Negative Control:
1. Start the test read by navigating to the “TEST” Menu, selecting the “QC Test”.  Select “Analyze Now” or “Incubate and Analyze”.  Follow the on-screen instructions. 

2. Remove the white tip cap from the bottom of the Aioprep and discard. 

3. While holding the Aioprep vertically, with both hands, squeeze near the bottom of the barrel to add 7 drops of Negative Control reagent (Sample Diluent) into the sample port of the test card.  

4. Discard the Aioprep device immediately.

5. Analyze Now: Incubate the test card at 19-27 C on the benchtop for 20 minutes.  Insert test card into the drawer of the Curian Analyzer and select “Curian HpSA QC Negative”.  Initiate read within 2 minutes after incubation is complete.  Analysis of the reaction must be initiated by the Curian Analyzer.  The analyzer will read the test card and automatically report the QC test result.

6. Incubate and Analyze: Insert test card immediately into the Curian Analyzer and select “Curian HpSA QC Negative”.  The Curian will time the incubation for 20 minutes then read the test card and automatically report the QC test result.

Positive Control:
1. Start the test read by navigating to the “TEST” Menu and select “QC Test”.  Select “Analyze Now” or “Incubate & Analyze”.  Follow the on-screen instructions

2. Invert the positive control bottle to mix. 

3. Remove the positive control tip cap and hold vertically to add 4 drops of positive control into the sample port of the test card.

4. Analyze Now: Incubate the test card at 19-27 C on the benchtop for 20 minutes.  Insert test card into the drawer of the Curian Analyzer and select “Curian HpSA QC Positive”.  Initiate read within 2 minutes after incubation is complete.  Analysis of the reaction must be initiated by the Curian Analyzer.  The Curian will read the test card and automatically report the QC test result.

5. Incubate and Analyze: If “Incubate & Analyze” is selected, insert the test card immediately into the Curian Analyzer and select “Curian HpSA QC Positive”.  The Curian Analyzer will time the incubation for 20 minutes then read the test card and automatically report the test result.

If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________                                                                                                   

Results of External QC and action(s) taken when control results are unacceptable are documented:________________________________________________________________

SPECIMEN:

Preferred sample types:

Unpreserved stool
Undesirable specimens:
1. Stool in transport media
2. Stool on swabs
3. Stool mixed with preservatives
Collection and Storage:
The patient specimen should be received in an airtight transport container and stored at 2-8 C until tested.  The specimen should be tested as soon as possible but may be held up to 72 hours at 2-8 C prior to testing.  If testing cannot be performed within this time frame, specimens should be frozen immediately upon receipt and stored frozen for up to 14 days at -20 C until tested.  Specimens may be frozen and thawed twice.  
TEST PROCEDURE:
Specimen Preparation and Test Procedure


     Note: Handle all waste and specimens as biohazardous. 
Note: Mix stool specimen thoroughly with the included wooden applicator (regardless of consistency) before testing.

1. Bring all test components, reagents and specimens to 19-27 C before testing.

2. Prior to sampling, mix stool specimen thoroughly with wooden applicator.  

3. Remove the top cap from the Aioprep device and set aside.
4. Using the Curian HpSA sample collection brush, insert the brush into the stool specimen to coat the bristles of the brush.  It is important that only the brush is covered with stool.  Do not get stool on the shaft or handle of the brush.  Refer to the examples below.

                  Correct 
                       Incorrect- Under-sampling
               Incorrect- Over-sampling

[image: image5.png]


[image: image6.png]


[image: image7.png]meridian BIOSCIENCE
S



 [image: image1.png]


                                               [image: image2.png]


                                                             [image: image3.png]



5. Insert the Curian HpSA sample collection brush into the Aioprep device, pushing the brush through the metering insert.  
6. Snap off the sample collection brush at the marked ‘Break’ point and discard the handle.  
7. Recap the Aioprep device, leaving the Curian HpSA sample collection brush in the device. 
8. Mix the sample by vortexing for approximately 5 seconds.  Do not invert or shake to mix. 

9. Use 1 Curian HpSA test card for each sample.  After mixing the sample, remove 1 test card from its foil pouch.  

10. Discard the pouch and desiccant.  Do not use test cards where the desiccant indicator has changed from blue to pink 
11. Label the test card with sample identification information.  Do not mark over or near the barcode on the test card
12. Remove the white tip cap from the bottom of the Aioprep and discard.  
13. While holding the Aioprep vertically, with both hands, squeeze near the bottom of the barrel to add 7 drops into the sample port of the test card.  

14. Discard the Aioprep device immediately.
15. Start the test read by navigating to the “TEST” Menu, selecting the appropriate test and following the on-screen instructions. 

16. Analyze Now: Incubate the test card at 19-27 C on the benchtop for 20 minutes.  Insert test card into the drawer of the Curian Analyzer and read test within 2 minutes of the end of incubation.  Analysis of the reaction must be initiated by the Curian Analyzer.  The analyzer will read the test card and automatically report the test result.

17. Incubate and Analyze: If “Incubate & Analyze” is selected, insert the test card immediately into the Curian Analyzer and the analyzer will time the incubation for 20 minutes then read the test card and automatically report the test result.

This facility’s procedure for specimen preparation is:____________________________  ______________________________________________________________________
This facility’s procedure for sample labeling is:_________________________________ ______________________________________________________________________
CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Result interpretation is completed automatically by the Curian Analyzer system.  The result will be shown on screen.  Results can be retrieved from instrument storage, printed and/or exported. 

Instrument Check: 

One of the following result interpretations will be generated by the Curian Analyzer for IC Check.

Pass: This indicates that the essential and critical components of the Curian Analyzer are working correctly. 

Fail: This indicates that there may be an issue with the Curian Analyzer.  Refer to the Curian operator’s manual if an error message has occurred.  Repeat testing to assist in trouble-shooting the error.  If repeat testing results in a failed output, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

External Quality Controls:

One of the following result interpretations will be generated by the Curian Analyzer for External Controls.

Pass: This indicates the test card and reagents are performing as intended.
Fail: This result indicates the test card and/ or reagents are not performing as intended or a user error has occurred.  Test should be repeated to assist in trouble-shooting the error.  If repeat testing results in a failed output, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor. 

Control Invalid/ Fail: This result indicates inadequate sample flow or incorrect sample preparation.  Repeat the test.  If repeat testing yields the same results, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance. 
Patient Test Results: 

One of the following result interpretations will be generated by the Curian Analyzer for patient specimens.

Positive: H. pylori stool antigen is detected in sample.
Negative: H. pylori stool antigen is not detected in sample.

Invalid: Result indicates inadequate flow of sample or incorrect sample preparation.  Repeat the test from the original stool specimen.  If repeat testing yields the same results, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:
Negative test:  Report test results as “H. pylori stool antigen is not detected in sample or    is below the level of detect for this assay.”


Positive test:  Report test result as “H. pylori stool antigen is detected in sample.”
EXPECTED VALUES:
Recent studies have shown that Helicobacter pylori infections are a major public health concern globally.  It is estimated that over half of the global population is infected with Helicobacter pylori.  In patients diagnosed with duodenal ulcers, however, it has been shown to be 90%-95%.  Currently, recommended eradication treatments have an overall eradication rate of 84.3% (82.1% to 86.4%).
The Curian HpSA assay detects the presence of H. pylori antigens in human stool.  Expected values for a given population should be determined for each laboratory.  The rate of positivity may vary depending on geographic location, method of specimen collection, handling and transportation, test employed and general health environment of patient population under study.  The observed prevalence of H. pylori in this study was 13.5% (73/542).  The prevalence of H. pylori by gender is provided below. 

	Gender
	Samples Positive by Comparator Method and Curian HpSA
	Prevalence in Tested Population

	Female
	42
	7.7%

	Male
	31
	5.7%

	Total
	73
	13.5%


ASSAY SPECIFICITY/CROSSREACTIVITY:

The specificity of Curian HpSA was tested utilizing the following bacterial, fungal and viral strains.  Each potentially cross-reactive microorganism was added at minimum concentrations of 1.0 x 107 CFU/mL (bacteria/fungi) or 1.0 x 105 TCID50/mL (for viruses) to a diluted, natural negative stool matrix and a contrived positive matrix sample.  No crossreactivity or microbial interference with the Curian HpSA assay was observed.

	Organism Name
	Strain ID
	Organism Name
	Strain ID

	Adenovirus 40
	Dugan
	Klebsiella pneumoniae
	ATCC 13883

	Aeromonas hydrophila
	ATCC 35654
	Proteus vulgaris
	CCUG 6380

	Bacillus cereus
	CCUG 52704
	Pseudomonas aeruginosa
	ATCC 39324

	Borrelia burgdorferi
	B31.5A19
	Rotavirus
	WA

	Campylobacter coli
	ATCC 10956
	Salmonella spp. Dublin
	ATCC 15480

	Campylobacter jejuni
	ATCC 29411
	Salmonella spp. Hilversum
	ATCC 15784

	Candida albicans
	ATCC 18804
	Salmonella spp. Minnesota
	ATCC 9700

	Citrobacter freundii
	ATCC 8090
	Salmonella typhimurium Group B
	ATCC 14028

	Clostridium difficile
	ATCC 43255
	Shigella boydii
	ATCC 9207

	Clostridium perfringens
	ATCC 12915
	Shigella dysenteriae
	ATCC 9361

	Enterobacter cloacae
	ATCC 15337
	Shigella flexneri
	ATCC 12022

	Enterococcus faecalis
	ATCC 49532
	Shigella sonnei
	ATCC 25931

	E. coli O157:H7
	ATCC 43895
	Staphylococcus aureus
	ATCC 6538

	E. coli
	ATCC 9637
	Staphylococcus aureus Cowan I
	ATCC 12598

	Escherichia fergusonii
	ATCC 35469
	Staphylococcus epidermidis
	ATCC 51625

	Haemophilus influenzae
	ATCC 9006
	Yersinia enterocolitica
	ATCC 23715


TEST FOR INTERFERING SUBSTANCES:
Interference testing was performed in the presence of chemical and biological substances introduced directly into contrived HpSA low positive and negative samples.  No interference was observed with the Curian HpSA assay for any of the substances tested.  The substances tested and concentrations evaluated are shown in the table below. 

	Substance (active ingredient(s))
	Test Concentration

	Barium Sulfate
	5% w/v (50 mg/mL)

	Benzalkonium chloride
	1% v/v

	Ciprofloxacin
	0.25% w/v (2.5 mg/mL)

	Ethanol
	1% v/v

	Hog gastric mucin
	3.5% w/v (35 mg/mL)

	Human blood (whole)
	40% v/v

	Human hemoglobin
	12.5% w/v (125 mg/mL)

	Human urine
	5% v/v

	Hydrocortisone
	1% w/v (10 mg/mL)

	Imodium® (Loperamide HCl, 1 mg/7.5 mL)
	5% v/v

	Kaopectate® (Bismuth subsalicylate 262 mg/15 mL)
	5% v/v

	Leukocytes
	0.05% v/v

	Mesalazine (5-Aminosalicylic acid)
	10% w/v (100 mg/mL)

	Metronidazole
	0.25% w/v (2.5 mg/mL)

	MiraLAX® (Polyethylene Glycol 3350, 17 g/dose)
	7% w/v (70 mg/mL)

	Mineral Oil
	10% v/v

	Mylanta® (per 10 mL: (Aluminum hydroxide 800 mg, Magnesium hydroxide 800 mg, Simethicone 80 mg)
	4.2 mg/mL (2.5% v/v)

	Naproxen Sodium
	5% w/v (50 mg/mL)

	Nonoxynol-9
	1% v/v

	Nystatin
	1% w/v (10 mg/mL)

	Palmitic acid (fecal fat)
	20% w/v (200 mg/mL)

	Pepto-Bismol® (Bismuth subsalicylate 525 mg/30 mL)
	5% v/v

	Phenylephrine
	1% w/v (10 mg/mL)

	Prilosec OTC® (Omeprazole 20 mg/tablet)
	5 mg/mL

	Sennosides
	1% w/v (10 mg/mL)

	Simethicone
	10% v/v

	Stearic acid (fecal fat)
	20% w/v (200 mg/mL)

	Tagamet HB 200® (Cimetidine 200 mg/tablet)
	5 mg/mL

	TUMS®
	5 mg/mL

	Vancomycin
	0.25% w/v (2.5 mg/mL)


LIMITATIONS OF THE PROCEDURE:
1. The Curian HpSA assay must only be used with the Curian Analyzer. 

2. Curian HpSA is a qualitative, in vitro diagnostic test.  The Curian Analyzer will only provide qualitative results.  This test is not intended to provide quantitative values.
3. The Curian HpSA assay is designed to be used with unpreserved stool samples.  DO NOT USE stool in transport media, on swabs, or mixed with preservatives.

4. Antimicrobials, proton pump inhibitors and bismuth preparations are known to suppress H. pylori and ingestion of these prior to H. pylori testing (culture, histology, rapid urease, UBT, antigen) may give a false negative result.  If a negative result is obtained for a patient ingesting these compounds within two weeks prior to performing the Curian HpSA assay, a false-negative result may be obtained.  The test should be repeated on a new specimen two weeks after discontinuing treatment. A positive result for a patient ingesting these compounds within two weeks prior to performing the Curian HpSA test, should be considered accurate. 

5. Performance characteristics have not been established in asymptomatic populations. 

6. A negative test result does not preclude the possibility of the presence of H. pylori antigen in the specimen which may occur if the level of antigen is below the detection limit of the test. 

7. No data exists on the effects of colonic washes, barium enemas, laxatives, or bowel preparations on the performance of the Curian HpSA assay. These procedures can result in extensive dilution or the presence of additives that may affect test performance.
PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience Curian HpSA
REFERENCES:

Refer to Directional Insert – Meridian Bioscience Curian HpSA
Meridian Bioscience

3471 River Hills Drive

Cincinnati, OH 45244 USA
Ph: (800) 343-3858, (513) 271-3700


SN11285 CLSI 07/20

