Procedure:  curian Shiga Toxin

Procedure and revision number: SN11289  REV. 04/23

Meridian Catalog:  760630

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.  Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The Curian Shiga Toxin assay, for use with the Curian Analyzer, is a rapid, qualitative, fluorescent immunoassay for the simultaneous detection and differentiation of Shiga toxin 1 (Stx1) and Shiga toxin 2 (Stx2) in a single test device.  It is intended for use with cultures derived from human stool specimens to aid in the diagnosis of disease caused by Shiga toxin producing Escherichia coli (STEC) infections.  Test results are to be used in conjunction with the patient’s clinical symptoms and history.

The Curian Shiga Toxin assay consists of a test strip enclosed in a plastic frame (test card), positive control reagent, and assay-specific Aioprep™ sample preparation device.  Curian Shiga Toxin is a lateral flow-based immunoassay for the simultaneous detection and differentiation of Shiga toxin 1 (Stx1) and Shiga toxin 2 (Stx2) in broth or plate cultures derived from human stool specimens.  Curian Shiga Toxin utilizes monoclonal antibodies specific to Stx1 and Stx2 as capture and detector antibodies, along with a polyclonal capture antibody.

The Aioprep is pre-filled with blue tinted Sample Diluent/Negative Control and contains a filter and a dropper tip.  A sample of the patient’s culture enriched stool specimen is transferred from the enrichment media to the Aioprep, using either the included transfer pipettes or a swab (not provided with the kit) to add the sample directly into the Sample Diluent/Negative Control.

The diluted sample is mixed and dispensed drop-wise into the sample port of the Curian Shiga Toxin test card.  If Stx1 and/or Stx2 antigens are present in the sample, they bind to the respective monoclonal detector antibodies conjugated to fluorescent particles, forming a complex.  As the complex moves through the test strip, the anti-Shiga toxin 1 and/or 2 monoclonal capture antibodies, each bound to the assay membrane at their respective Stx1 and Stx2 test positions of the strip, bind the complexes and yield the respective test lines.  When antigen is not present, a complex is not formed, and test lines do not form.  As the sample continues to move further up the test strip, the polyclonal capture antibody, bound to the assay membrane at the control position of the strip, binds the conjugated antibodies and yields the control line.  A line at the control position of the test strip should be present each time a sample or external control is tested.  If the control line is not generated, adequate sample flow has not occurred, and the Curian Analyzer will consider the test invalid.

SPECIMEN:

Preferred sample types:

- unpreserved stool samples

- stools preserved in Cary-Blair or C&S media

Undesirable specimens:

- stool on swabs or in other transport media (e.g., formalin, polyvinyl alcohol)

- colonic washes

- barium enemas

- stool containing laxatives

- bowel preparations

Specimen Collection, Handling and Storage:

This procedure is designed to be used with cultures derived from unpreserved stool or stool preserved in Cary-Blair or C&S transport media.

- Stool specimen storage and handling for the culture of STEC organisms: The patient stool specimen should be received in an airtight container and either be frozen (≤ -20 C) or placed at 2-8 C immediately after collection. The refrigerated unpreserved specimens should be cultured within two (2) hours of collection. If culturing cannot be performed within 2 hours, the specimen should be placed in a Cary-Blair or C&S transport media. Samples in Cary-Blair or C&S transport media may be stored at 2-8 C for up to five (5) days prior to preparing cultures. If culturing cannot be performed within this time, the specimens may be frozen at ≤ -20 C for up to fourteen (14) days prior to preparing cultures.

| Storage Temperatures and Maximum Storage Times for Stool Specimens |
| --- |
| Specimen Conditions | Storage Temperature | Maximum Storage Time |
| Unpreserved specimens | 2 – 8 C (refrigerated) | 2 hours |
|  | ≤ -20 C (frozen) | 14 days |
| Preserved specimens in Cary-Blair or C&S media | 2 – 8 C (refrigerated) | 5 days |
|  | ≤ -20 C (frozen) | 14 days |

- Storage of broth showing growth prior to preparation in Aioprep / Sample Diluent for Curian Shiga Toxin testing: Broths with growth may be held for up to seven (7) days at 2-8 C before testing with Curian Shiga Toxin. If testing is not performed within this time period, the broth should be frozen at ≤ -20 C for up to twenty-one (21) days. Broths may be frozen and thawed twice.

| Storage Temperatures and Maximum Storage Times for Broth Cultures |
| --- |
| Broth cultures with growth | Storage Temperature | Maximum Storage Time |
|  | 2 – 8 C (refrigerated) | 7 days |
|  | ≤ -20 C (frozen) | 21 days |

- Storage of processed cultures in Aioprep / Sample Diluent prior to Curian Shiga Toxin testing: Broths with growth that have been prepared in the Aioprep/ Sample Diluent may be held at room temperature (19 - 27 C) for up to eight (8) hours or at refrigerated temperature (2-8 C) for up to twenty-four (24) hours prior to testing with the Curian Shiga Toxin assay. SMAC agar plates with growth prepared in the Aioprep / Sample Diluent may be held at room temperature (19 - 27 C) for up to two (2) hours or at refrigerated temperature (2 - 8 C) for up to twenty-four (24) hours prior to testing with the Curian Shiga Toxin assay.

| Storage Temperatures and Maximum Storage Times for Processed Cultures in Aioprep / Sample Diluent |
| --- |
| Specimen Conditions | Storage Temperature | Maximum Storage Time |
| Prepared broth cultures | 19 - 27 C (room temperature) | 8 hours |
|  | 2 - 8 C (refrigerated) | 24 hours |
| Prepared SMAC agar plate cultures | 19 - 27 C (room temperature) | 2 hours |
|  | 2 - 8 C (refrigerated) | 24 hours |

Specimen Preparation and Enrichment:

Note: Power-on the Curian Analyzer and complete Instrument Check (IC) test, and External Quality Testing before starting the sample preparation and test procedure.

Transfer pipette (PIPETTE I; 50 µL and 175 µL) for use in SPECIMEN PREPARATION / ENRICHMENT steps of this package insert:

## BROTH METHOD:

  - Thoroughly mix stool specimen by vortexing for approximately 5 seconds regardless of consistency.

  - Using a pipettor or transfer pipette (PIPETTE I, supplied with the kit):

    - Human stool specimens, unpreserved: Add 50 µL of unpreserved specimen (first mark from tip of pipette; refer to image below) to a culture tube containing 8 mL of GN broth or 5 mL of MacConkey broth.

    - Human stool specimens, preserved in Cary-Blair or C&S transport media: Add 175 µL of the preserved specimen (second mark from tip of pipette; refer to image below) to a culture tube containing 8 mL GN broth or 5 mL MacConkey broth.

Unpreserved stool – use first mark from tip.

Preserved stool – use second mark from pipette tip.

  - Vortex inoculated broth for approximately 5 seconds.

  - Incubate inoculated broth with broth tube cap loose at 35-39 C for 16-24 hours.

  - Examine the broth tubes for growth. DO NOT PROCEED WITH TESTING if the broth tube does not exhibit growth after incubation as falsely negative results may occur. Repeat the broth enrichment using the same stool sample or with a new sample collected from the patient. If the original broth enrichment was performed with GN broth, GN broth can be used in the second attempt or alternatively, MacConkey broth can be used instead and vice versa. The sample can also be recultured using the SMAC plate method. If agar cultures are used, proceed with instructions provided in the SMAC PLATE METHOD BELOW. Use only broth cultures that exhibit growth in the following TEST PROCEDURE steps.

## SMAC AGAR PLATE METHOD:

- Thoroughly mix stool specimen by vortexing for approximately 5 seconds regardless of consistency.

- Use a polyester swab to inoculate stool samples onto SMAC agar plates without tellurite or cefixime. (NOTE: Tellurite and cefixime inhibit growth of non-O157:H7 E. coli.)

- Incubate the inoculated plate for 18-24 hours at 35-39 C.

- Examine the agar plates for growth. DO NOT PROCEED WITH TESTING if the agar plate does not exhibit growth after incubation as falsely negative results may occur. Repeat the agar enrichment using the same stool sample or with a new sample collected from the patient. The sample can also be recultured using the broth method. If broth cultures are used, proceed with instructions provided in the BROTH METHOD ABOVE. Use only agar cultures that exhibit growth in the following TEST PROCEDURE steps.

MATERIALS AND EQUIPMENT:

Materials Provided:

### The maximum number of tests obtained from this test kit is listed on the outer box.

- Curian Shiga Toxin Test Card: A test strip enclosed in a plastic frame which is in a foil pouch with a desiccant. Supplied ready to use.

- Curian Shiga Toxin Aioprep Sample Preparation Device/ Negative Control: A buffered aqueous protein solution containing blue dye and 0.094% sodium azide. The Aioprep device is fitted with a turquoise top cap and a dropper tip with a white cap. Supplied ready to use.

- Curian Shiga Toxin Positive Control: Inactivated Shiga toxins 1 and 2 in an aqueous phosphate buffered solution containing 0.094% sodium azide. Supplied ready to use.

- Pipette I: Transfer pipettes graduated to 50 µL and 175 µL.

- Pipette II: Transfer pipettes graduated up to 1.0 mL.

Materials Not Provided:

## All Methods:

- Disposable gloves, powder free

- Pipettor and tips

## Broth Method:

- Gram Negative (GN) or MacConkey broth

- Culture tubes

## SMAC Agar Plate Method:

- Sorbitol-MacConkey agar plate without tellurite or cefixime

- Polymyxin B (20 mg/mL in water)

- Polyester swabs

Equipment Not Provided:

- Vortex mixer

- Interval timer (Optional)

- Incubators, 35-39 C

- Curian Analyzer System, Meridian Bioscience, Inc. Catalog 610190

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- Store the kit at the temperature indicated on labeling when not in use.

- Follow Good Laboratory Practices10 (GLPs) when handling the kit, reagents and components.

- Do not interchange transfer pipettes (PIPETTE I and PIPETTE II) or Aiopreps between assays. Specimens must be sampled and prepared with the transfer pipettes (PIPETTE I and PIPETTE II) and used with the Curian Shiga Toxin Aioprep device provided with the kit. These components are not interchangeable with other assays.

- Handle and dispose of all human specimens as if they are biologically hazardous.

- Inspect foil pouch before removing the test card. Do not use test cards that have holes in the foil pouch or where the pouch has not been completely sealed.

- Do not use test cards where the desiccant indicator has changed from blue to pink.

- Do not interchange positive control or test cards between kit lots.

- Do not mark over or near the barcode on the test card.

- Sample preparation in the Aioprep can be performed inside a laminar flow hood. The Curian Shiga Toxin test card must be incubated outside of a laminar flow hood.

STORAGE REQUIRMENTS:

1.    The expiration date is indicated on the kit box label.

- Store the kit at 2 - 8 C as indicated on the labeling.

- Return all kit components to the indicated storage temperature after use.

At this facility, kits are stored: _______________________________________________

CALIBRATION:

Not applicable to this assay.

# OPERATION OF THE CURIAN ANALYZER SYSTEM:

The Curian Analyzer is an easy to use, menu driven analyzer system. Instructions to complete testing are provided on the analyzer touchscreen and in the Curian Analyzer operator manual.

- Power ON the Curian Analyzer by pressing the power button located on the left side of the analyzer. The Curian Analyzer will initialize and perform Self Test.

- Instrument Checks, Quality Control testing, and Patient Specimen testing are completed by navigating from the Home screen to the test menus. User ID and Sample ID entries are required for Quality Control and specimen testing.

- Complete use and operation of the Curian Analyzer System by following the instructions in the Curian Analyzer operator manual and on-screen prompts. Refer to operator manual for detailed instructions for the Curian Analyzer.

QUALITY CONTROL:

Note: Good Laboratory Practice (GLP) guidelines recommend the use of control material. Users should follow the appropriate federal, state, and local guidelines concerning the running of external quality controls.

Bring all test components and reagents to 19 - 27 C before testing. Use one Curian test card for each external QC test (Negative and Positive Controls). When ready to perform testing, remove the test card from its foil pouch. Discard the pouch and desiccant. Do not use if desiccant is pink. Do not mark over the barcode on the test card.

Start the Test Read by navigating to the ‘TEST’ Menu, selecting the ‘QC TEST’, and then selecting ‘ANALYZE NOW’ or ‘INCUBATE AND ANALYZE’. Follow the on-screen instructions.

Negative Control:

- Remove the top turquoise cap from Aioprep device completely, then recap the Aioprep tightly. Remove the white tip cap from the bottom of the Aioprep and discard. While holding the Aioprep vertically, squeeze in the center of the barrel to add 3 free-falling drops of Negative Control reagent (Sample Diluent) into the ‘SAMPLE’ port of the test card. Discard the Aioprep device immediately.

## For Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes.

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be  initiated by the Curian Analyzer.

  - Select ‘Curian Shiga Toxin QC Negative’.

  - The Curian Analyzer will analyze the test card and automatically report the test result.

## For Incubate and Analyze:

- Enter Sample ID and press Ok.

- Insert test card immediately into the Curian Analyzer and close the drawer.

- Select ‘Curian Shiga Toxin QC Negative’.

- The Curian Analyzer will time the incubation for 20 minutes, then read the test card and automatically report the test result.

Positive Control:

- Invert the positive control bottle to mix. Remove the Positive Control tip cap. Holding vertically, squeeze 3 drops of Positive Control into the ‘SAMPLE’ port of the test card.

## For Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes.

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be initiated by the Curian Analyzer.

  - Select ‘Curian Shiga Toxin QC Positive’.

  - The Curian Analyzer will analyze the test card and automatically report the test result.

## For Incubate and Analyze:

- Enter Sample ID and press Ok.

- Insert the test card immediately into the drawer of the Curian Analyzer and close the drawer.

- Select ‘Curian Shiga Toxin QC Positive’.

- The Curian Analyzer will time the incubation for 20 minutes, then read the test card and automatically report the test result.

If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor.

INSTRUMENT CHECK (IC) TEST PROCEDURE:

Note: The analyzer’s External Control (IC Test) has a default schedule of 30 days. The IC Test will be required to be performed prior to running a QC test or patient test if this has not been performed within this timeframe. The user will be locked-out of running a QC test or patient test if this test is overdue.

- Power ON the Curian by pressing the power button located on the side of the Curian. The analyzer will initialize and perform SELF TEST (automatic) checks.

- From the HOME screen, select ‘TEST’, and then select ‘INSTRUMENT CHECK’.

- Insert the Fluorescent IC Card into the drawer. A ‘PASS’ or ‘FAIL’ result will be displayed once the analysis is complete.

If the expected result is not observed, repeat the IC Test as the first step in determining the root cause of the failure. Refer to the Curian Operator’s Manual for detailed instructions for the analyzer.

### This test should be performed per applicable local, state, or federal regulations or accrediting agencies.

There are four controls for the Curian Campy test system. There is an external and internal control for both the analyzer and assay.

- Analyzer Internal Control: Self Test (automatic)

- Analyzer External Control: Instrument Check (IC) Test

- Assay External Controls: Positive and Negative Controls (QC test)

- Assay Internal Control: Control Line

QC Testing Frequency and Documentation:

For this facility, External QC is run: _________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented: ________________________________________________________________

TEST PROCEDURE:

Bring all test components, reagents, and samples to room temperature (19-27 C) before testing. Use one Curian Shiga Toxin test card for each sample. When ready to perform testing, inspect the pouch for damage. If undamaged, remove test card from its foil pouch. Discard the pouch and desiccant. Do not use if desiccant is pink. Do not cover or mark on the barcode.

Transfer pipette (PIPETTE II; 1.0 mL) for use in TEST PROCEDURE steps of this package insert:

## BROTH METHOD:

  - Use only broth cultures that exhibit growth. Mix broth culture thoroughly by vortexing broth tube for approximately 5 seconds.

  - Remove the turquoise top cap of the Aioprep device.

## Note: Prepared broth cultures in Aioprep / Sample Diluent may be held at room temperature (19 - 27 C) for up to 8 hours or at refrigerated temperature (2 - 8 C) for up to 24 hours prior to proceeding to step 2 for testing.

  - Use the provided transfer pipette (PIPETTE II) to transfer 1 mL of cultured broth (fourth mark from tip of pipette) into the Aioprep device. Ensure that the specimen is completely expelled from the pipette into the Aioprep.

  - Do not invert or shake Aioprep. Recap the Aioprep tightly and vortex to mix for approximately 5 seconds.

Cultured broth – use fourth mark from pipette tip.

## SMAC AGAR PLATE METHOD:

- Use only agar cultures that exhibit growth.

- Remove the turquoise top cap of the Aioprep device.

## Note: Prepared SMAC agar plate cultures in Aioprep / Sample Diluent may be held at room temperature (19 - 27 C) for up to 2 hours or at refrigerated temperature (2 - 8 C) for up to 24 hours prior to proceeding to step 2 for testing.

- Use a pipettor to add 5 µL of Polymyxin B to the Aioprep device.

- Using a polyester swab, sweep a few times across the confluent growth area of the plate or several isolated colonies, avoiding mucoid colonies. (NOTE: Mucoid colonies may interfere with migration of the sample.)

- Swirl swab carrying the colony sweep in the Polymyxin B/Sample Diluent Solution and rotate the swab for 5 seconds to enhance the release of Shiga toxins from the organisms.

- Do not invert or shake Aioprep. Recap the Aioprep tightly, vortex for approximately 5 seconds, and incubate for 30 minutes at 35-39 C.

- Vortex the Polymyxin B/Diluted Sample for approximately 5 seconds and allow to acclimate to 19-27 C.

- Remove the white tip cap from the bottom of the Aioprep and discard. While holding the Aioprep vertically, squeeze in the center of the barrel to add 3 free- falling drops into the ‘SAMPLE’ port of the test card. (NOTE: Failure to add 3 free-falling drops may lead to invalid test results.) Discard the Aioprep device immediately.

- Start the Test Read by navigating to the ‘TEST’ Menu, selecting ‘Analyze Now’ or ‘Incubate and Analyze’, and then following the on-screen instructions.

## Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes. (NOTE: Over incubation of tests may lead to false-positive or invalid test results.)

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be initiated by the Curian Analyzer.

  - The Curian will analyze the test card and automatically report the test result.

## Incubate and Analyze:

- Enter Sample ID and press Ok.

- Immediately insert the test card into the drawer of the Curian Analyzer and close the drawer. The analyzer will time the incubation for 20 minutes.

- After the end of the incubation, the Curian will analyze the test card and automatically report the test result.

This facility’s procedure for specimen preparation is:____________________________  ______________________________________________________________________

This facility’s procedure for sample labeling is:_________________________________ ______________________________________________________________________

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Result interpretation is completed automatically by the Curian Analyzer system. The result will be shown on screen. Results can be retrieved from analyzer storage, printed and/or exported.

Patient Test Results:

One of the following result interpretations will be generated by the Curian Analyzer.

| Curian Analyzer Result | Result Interpretation |
| --- | --- |
| Stx 1: Positive Stx 2: Negative | Shiga toxin 1 is present in sample. |
| Stx 1: Negative Stx 2: Positive | Shiga toxin 2 is present in sample. |
| Stx 1: Positive Stx 2: Positive | Shiga toxin 1 and Shiga toxin 2 are both present in sample. |
| Stx 1: Negative Stx 2: Negative | Shiga toxin 1 and Shiga toxin 2 are not present in sample. |
| Invalid | Result indicates inadequate flow of sample. Repeat the test. If repeat testing yields the same results, contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance. |

External Quality Controls:

One of the following result interpretations will be generated by the Curian.

- Pass: This indicates the test card and reagents are performing as intended.

- Fail: This result indicates the test card and/ or reagents are not performing as intended or a user error has occurred. Test should be repeated to assist in trouble-shooting the error. If repeat testing results in a failed output, please contact your Meridian Bioscience Representative.

- Control Invalid: This result indicates inadequate flow of sample. Repeat the test. If repeat testing yields the same results, contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance.

Instrument Check:

One of the following result interpretations will be generated by the Curian Analyzer.

- Pass: This indicates that the essential and critical components of the Curian Analyzer are working correctly.

- Fail: This indicates that there may be an issue with the Curian Analyzer. Refer to the Curian Operator Manual if an error message has occurred. Repeat testing to assist in trouble-shooting the error. If repeat testing results in a failed output, please contact Meridian’s Technical Services Department at 1-800-343- 3858 (US) or your local distributor for additional assistance.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Shiga Toxin 1 and Shiga toxin 2 not detected, or below the level of detect in sample.

Positive Shiga Toxin 1:  Shiga toxin 1 detected, Shiga toxin 2 not detected, or below the level of detect in sample.

Positive Shiga Toxin 2:  Shiga toxin 1 not detected, or below the level of detect, Shiga toxin 2 detected in sample.

Positive Shiga toxin 1 and Shiga toxin 2:  Shiga toxin 1 and Shiga toxin 2 detected in sample.

EXPECTED VALUES:

The prevalence of Shiga toxin-producing E. coli (STEC) infection observed during the 2022 study period for prospectively collected specimens was 0.3% for both Shiga toxin 1 (5/1538) and Shiga toxin 2 (4/1538). The prevalence of STEC infection by sex assigned at birth for the tested population was 0.3% for Stx1 and 0.2% for Stx2 in males and 0.3% for both Stx1 and Stx2 in females. The age demographic for the tested population ranged from less than 2 years to 60 years and older, with the mean age of 43 years old.

# ANALYTICAL SENSITIVITY:

Analytical sensitivity studies were performed to determine the analytical limit of detection (LoD) of quantified Shiga toxin 1 (Stx1) and Shiga toxin 2 (Stx2) in cultures derived from human stool specimens for the Curian Shiga Toxin assay. The LoD is defined as the lowest concentration of the target analyte that produces positive results ≥ 95% of the time.

ASSAY SPECIFICITY:

CROSSREACTIVITY / MICROBIAL INTERFERENCE:

The Curian Shiga Toxin assay was evaluated for cross-reactivity and microbial interference with the organisms listed below. Unless otherwise indicated, each organism was tested at minimum concentrations of 1.0 x 107 CFU/mL for bacteria/fungi or 1.0 x 105 TCID50/mL for viruses. None of the organisms showed cross- reactivity or microbial interference in the Curian Shiga Toxin assay.

| Aeromonas hydrophila | Gardnerella vaginalis |
| --- | --- |
| Bacillus subtilis | Helicobacter pylori |
| Bacteroides fragilis | Klebsiella pneumoniae |
| Campylobacter coli | Lactobacillus acidophilus |
| Campylobacter concisus | Proteus vulgaris |
| Campylobacter fetus | Providencia stuartii |
| Campylobacter hyointestinalis | Pseudomonas aeruginosa |
| Campylobacter jejuni | Pseudomonas fluorescens |
| Candida albicans | Salmonella enterica susb. enterica serovar Hilversum |
| Citrobacter freundii | Salmonella enterica susb. enterica serovar Typhimurium |
| Clostridium difficile | Salmonella minnesota |
| Clostridium perfringens | Serratia liquefaciens |
| Enterobacter cloacae | Serratia marcescens |
| Enterococcus faecalis | Shigella boydii |
| Escherichia coli (non-toxigenic) | Shigella flexneri |
| Escherichia coli EIEC | Shigella sonnei |
| Escherichia coli EPEC | Staphylococcus aureus |
| Escherichia coli ETEC | Staphylococcus aureus (Cowan’s) |
| Escherichia coli O157:H7 (non-toxigenic) | Staphylococcus epidermidis |
| Escherichia fergusonii | Streptococcus equisimilis subsp. dysgalactiae |
| Escherichia hermanii | Yersinia enterocolitica |
| Human Adenovirus 2 | Human Coxsackievirus B1 |
| Human Adenovirus 14 | Human Enterovirus 69 |
| Human Adenovirus 40 | Human Rotavirus |
| Human Adenovirus 41 | Feline calicivirus |
| Human Coxsackievirus A9 |  |

Shigella species shown to be reactive with the Curian Shiga Toxin assay:

S. dysenteriae (strain ATCC 9361) was found to be Shiga toxin 1 (Stx1) positive at concentrations greater than 1.25 x 106 CFU/mL in the Curian Shiga Toxin assay.

TEST FOR INTERFERING SUBSTANCES:

The chemical and biological substances listed below were evaluated at the indicated concentrations for interference in the Curian Shiga Toxin assay. None of the substances showed interference with the Curian Shiga Toxin assay performance.

Barium Sulfate (5% w/v)

Ciprofloxacin (0.25% w/v)

Hog gastric mucin (3.5% w/v)

Human blood (40% v/v)

Human urine (5% v/v)

Imodium® A-D (5% v/v)

Kaopectate® (5% v/v)

Leukocytes (0.05% v/v)

Mylanta® (8.400 mg/mL)

Palmitic Acid/Fecal Fat (40% w/v)

Pepto-Bismol® (5% v/v)

Prilosec OTC® (5 µg/mL)

Stearic Acid/Fecal Fat (40% w/v)

Tagamet® (5 µg/mL)

TUMS® (50 µg/mL)

Naproxen sodium (0.5% w/v)

Metronidazole (0.25% w/v)

Vancomycin (0.25% w/v)

LIMITATIONS OF THE PROCEDURE:

- The Curian Shiga Toxin assay must only be used with the Curian Analyzer.

- Curian Shiga Toxin is a qualitative, in vitro diagnostic test. The Curian Analyzer will only provide qualitative results. This test is not intended to provide quantitative results.

## PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience curian Shiga Toxin

## REFERENCES:

Refer to Directional Insert – Meridian Bioscience curian Shiga Toxin