Procedure:  curian CAMPY

Procedure and revision number: SN11287  REV. 01/22

Meridian Catalog:  760730

Institution:

Address:

Department:

| Prepared By: | Date Adopted | Supersedes procedure #: |
| --- | --- | --- |

| Distributed To: | # of Copies | Distributed To: | # of Copies |
| --- | --- | --- | --- |

NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.  Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

Curian Campy, for use with the Curian Analyzer, is a rapid, qualitative fluorescent immunoassay for the detection of a Campylobacter-specific antigen in human fecal specimens. Curian Campy is intended to detect C. jejuni, C. coli, C. upsaliensis, and C. lari in human stool from patients with signs and symptoms of gastroenteritis. The test is intended for use with unpreserved fecal specimens or preserved fecal specimens in transport media. Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures. Curian Campy is intended to aid in the diagnosis of Campylobacter infection.

The Curian Campy assay consists of a test strip enclosed in a plastic frame (test card), positive control reagent, and assay-specific Aioprep™ sample preparation device. Curian Campy is a lateral flow-based immunoassay for the direct detection of a Campylobacter-specific antigen in human stool. Curian Campy utilizes monoclonal antibodies, specific for an antigen common to C.jejuni, C.coli, C. lari and C. upsaliensis, as the capture and detector antibodies.

The Aioprep is pre-filled with blue tinted Sample Diluent/Negative Control and contains a filter and a sample metering device. A sample of the patient’s stool specimen is transferred from the collection container to the Aioprep, using either the included transfer pipette (for preserved or unpreserved liquid/semi-solid specimens), or the assay-specific sample collection brush (for preserved or unpreserved formed/solid specimens) available separately as an accessory. When using the sample collection brush, the brush with sample is inserted through the metering device to remove excess stool and then pushed directly into the Sample Diluent/Negative Control. When using the transfer pipette, the sample is added directly into the Sample Diluent/Negative Control.

The diluted sample is mixed and dispensed drop-wise into the sample port of the Curian Campy test card. If present, the Campylobacter antigen binds to the monoclonal detector antibody conjugated to fluorescent particles, forming a complex. As the sample moves through the test strip, the anti-Campylobacter capture antibody, bound to the assay membrane at the test position of the strip, binds the complex and yields a test line. When antigen is not present, a complex is not formed, and a test line will not form. As the sample continues to move further up the test strip, the polyclonal capture antibody, bound to the assay membrane at the control position of the strip, binds the conjugated antibody and yields a control line. A line at the control position of the test strip should be present each time a sample or external control is tested. If the control line is not generated, adequate sample flow has not occurred, and the Curian Analyzer will consider the test invalid.

SPECIMEN:

Preferred sample types:

- unpreserved stool samples

- stools preserved in Cary-Blair or C&S media

Undesirable specimens:

- stool on swabs or in other transport media (e.g., formalin, polyvinyl alcohol)

- colonic washes

- barium enemas

- stool containing laxatives

- bowel preparations

Collection and Preparation:

This procedure is designed to be used with unpreserved stool or stool preserved in Cary-Blair or C&S transport media. DO NOT USE stool on swabs or mixed with fixatives. The patient specimen should be received in an airtight transport container and stored at 2 - 8 C until tested. The specimen should be tested as soon as possible, but preserved specimens may be held for up to 96 hours (four days) at room temperature (19 - 27 C), or unpreserved and preserved specimens (in Meridian Para-Pak C&S media) may be held for up to 7 days at 2 - 8 C prior to testing. If testing cannot be performed within this time frame, unpreserved specimens or specimens preserved in Meridian Para-Pak C&S media should be frozen immediately upon receipt and stored frozen for up to 59 days (or eight weeks) at ≤ -20 C until tested. Specimens may be frozen and thawed up to 5 times.

| Specimen Condition | Storage Temperature | Maximum Storage Time |
| --- | --- | --- |
| Unpreserved specimens | 2 – 8 C (refrigerated) | 7 days |
|  | ≤ -20 C (frozen) | 8 weeks |
| Preserved specimens in Cary-Blair or C&S media | 19 – 27 C (room temperature) | 96 hours |
|  | 2 – 8 C (refrigerated)1 | 7 days |
|  | ≤ -20 C (frozen)2 | 8 weeks |
| Prepared specimens in Aioprep / Sample Diluent | 19 – 27 C (room temperature) | 8 hours |
|  | 2 – 8 C (refrigerated) | 24 hours |

1 Refrigerated storage validated with Meridian Para-Pak C&S media only. If other C&S or Cary-Blair transport media are used, specimens should be stored according to the recommendations in the respective transport media package insert.

2 Frozen storage validated with Meridian Para-Pak C&S media only. If other C&S or Cary-Blair transport media are used, specimens should be stored according to the recommendations in the respective transport media package insert.

MATERIALS AND EQUIPMENT:

Materials Provided:

The maximum number of tests obtained from this test kit is listed on the outer box.

- Curian Campy Test Card: A test strip enclosed in a plastic frame which is in a foil pouch with a desiccant. Supplied ready to use.

- Curian Campy Aioprep Sample Preparation Device/ Negative Control: The Aioprep device is fitted with a metering insert and dropper tip. Supplied ready to use. A buffered protein solution containing blue dye and 0.094% sodium azide.

- Curian Campy Positive Control: Inactivated Campylobacter jejuni antigen in a phosphate buffered solution containing 0.094% sodium azide. Supplied ready to use.

- Transfer Pipettes

Materials Provided Separately:

- Curian Campy Stool Collection Brushes, Meridian Bioscience, Inc. Catalog 11516 (To order brushes, please contact Meridian’s Technical Services Department at 1-800-343-3858 or your local distributor.)

Materials Not Provided:

- Disposable gloves, powder free.

- Wooden applicator sticks

- Transport Media (Cary Blair or C&S)

Equipment Not Provided:

- Vortex mixer

- Interval timer (Optional)

- Curian Analyzer System, Meridian Bioscience, Inc. Catalog 610190

PERFORMANCE CONSIDERATIONS:

- All reagents are for in vitro diagnostic use only.

- Rx Only

- Store the kit at the temperature indicated on labeling when not in use.

- Do not interchange transfer pipettes, sample collection brushes, or Aiopreps between assays. Specimens must be sampled and prepared with the transfer pipettes or the Curian Campy sample collection brushes (provided as a separate accessory item) and used with the Curian Campy Aioprep device. These components are not interchangeable with other assays.

- Handle and dispose of all human specimens as if they are biologically hazardous.

- Inspect foil pouch before removing the test card. Do not use test cards that have holes in the foil pouch or where the pouch has not been completely sealed.

- Do not use test cards where the desiccant indicator has changed from blue to pink.

- Inspect Curian Campy Aioprep Sample Preparation Device/ Negative Control prior to use. Tap the Aioprep to ensure the liquid is in the main chamber of the Aioprep.

- Do not interchange positive control, Aiopreps, or test cards between kit lots.

- Do not mark over or near the barcode on the test card.

- The Curian Campy test card must be incubated outside of a laminar flow hood; this can be done either inside the analyzer or on the lab bench. Sample preparation in the Aioprep can be performed inside a laminar flow hood.

- Do not use transfer pipettes with formed and solid stool specimens.

STORAGE REQUIRMENTS:

1.    The expiration date is indicated on the kit box label.

- Store the kit at 2 - 8 C as indicated on the labeling.

- Return all kit components to the indicated storage temperature after use.

At this facility, kits are stored:_______________________________________________

CALIBRATION:

Not applicable to this assay.

QUALITY CONTROL:

Note: Good Laboratory Practice (GLP) guidelines9 recommend the use of control material. Users should follow the appropriate federal, state, and local guidelines concerning the running of external quality controls.

Bring all test components and reagents to 19 - 27 C before testing. Use one Curian Campy test card for each external QC test (Negative and Positive Controls). When ready to perform testing, remove the test card from its foil pouch. Discard the pouch and desiccant. Do not use if desiccant is pink. Do not mark over the barcode on the test card.

Start the Test Read by navigating to the ‘TEST’ Menu, selecting the ‘QC TEST’, and then selecting ‘ANALYZE NOW’ or ‘INCUBATE AND ANALYZE’. Follow the on-screen instructions.

Negative Control:

- Remove the white tip cap from the bottom of the Aioprep and discard. While holding the Aioprep vertically, squeeze in the center of the barrel to add 3 drops of Negative Control reagent (Sample Diluent) into the ‘SAMPLE’ port of the test card. Discard the Aioprep device immediately.

## For Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes.

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be initiated by the Curian Analyzer.

  - Select ‘Curian Campy QC Negative’.

  - The Curian Analyzer will analyze the test card and automatically report the test result.

## For Incubate and Analyze:

- Enter Sample ID and press Ok.

- Insert test card immediately into the Curian Analyzer.

- Select ‘Curian Campy QC Negative’.

- The Curian Analyzer will time the incubation for 20 minutes, then read the test card and automatically report the test result.

Positive Control:

- Invert the positive control bottle to mix. Remove the Positive Control tip cap. Holding vertically, squeeze 3 drops of Positive Control into the ‘SAMPLE’ port of the test card.

## For Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes.

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be initiated by the Curian Analyzer.

  - Select ‘Curian Campy QC Positive’.

  - The Curian Analyzer will analyze the test card and automatically report the test result.

## For Incubate and Analyze:

- Enter Sample ID and press Ok.

- Insert the test card immediately into the drawer of the Curian Analyzer and close the drawer.

- Select ‘Curian Campy QC Positive’.

- The Curian Analyzer will time the incubation for 20 minutes, then read the test card and automatically report the test result.

If the expected control reactions are not observed, repeat the control tests as the first step in determining the root cause of the failure. If control failures are repeated, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor. Do NOT perform patient testing if either of the QC test results are incorrect.

INSTRUMENT CHECK (IC) TEST PROCEDURE:

Note: The analyzer’s External Control (IC Test) has a default schedule of 30 days. The IC Test will be required to be performed prior to running a QC test or patient test if this has not been performed within this timeframe. The user will be locked-out of running a QC test or patient test if this test is overdue.

Power ON the Curian Analyzer by pressing the power button located on the side of the Curian Analyzer. The analyzer will initialize and perform SELF TEST (automatic) checks.

From the HOME screen, select ‘TEST’, and then select ‘INSTRUMENT CHECK’.

Insert the Fluorescent IC Card into the drawer. A ‘PASS’ or ‘FAIL’ result will be displayed once the analysis is complete.

If the expected result is not observed, repeat the IC Test as the first step in determining the root cause of the failure. Refer to the Curian Analyzer Operator’s Manual for detailed instructions for the analyzer.

There are four controls for the Curian Campy test system. There is an external and internal control for both the analyzer and assay.

- Analyzer Internal Control: Self Test (automatic)

- Analyzer External Control: Instrument Check (IC) Test

- Assay External Controls: Positive and Negative Controls (QC test)

- Assay Internal Control: Control Line

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________

Results of External QC and action(s) taken when control results are unacceptable are documented:________________________________________________________________

TEST PROCEDURE:

PROCEDURAL NOTES

The Curian Campy transfer pipette is diagrammed below:

OPERATION OF THE CURIAN ANALYZER SYSTEM

The Curian Analyzer is an easy to use, menu driven analyzer system.  Instructions to complete testing are provided on the analyzer touchscreen and in the Curian Operator’s Manual.

- Power ON the Curian Analyzer by pressing the power button located on the left side of the analyzer. The Curian Analyzer will initialize and perform SELF TEST.

- Instrument Checks, Quality Control testing, and Patient Specimen testing are completed by navigating from the Home screen to the test menus. User ID and Sample ID entries are required for Quality Control and specimen testing.

- Complete use and operation of the Curian Analyzer System by following the instructions in the Curian Operator’s Manual and on-screen prompts. Refer to operator’s manual for detailed instructions for the Curian Analyzer.

Specimen Preparation / Test Procedure:

## Note: Handle all waste and specimens as biohazardous.

Note: Before testing, mix stool specimen thoroughly by vortexing or inverting for liquid/semi-solid specimens, or with a wooden applicator stick for formed/solid specimens.

Bring all test components, reagents and specimens to 19 - 27 C before testing. Use one Curian Campy test card for each sample. When ready to perform testing, inspect the pouch for damage. If undamaged, remove test card from its foil pouch. Discard the pouch and desiccant. Do not use if desiccant is pink. Do not cover or mark on the barcode.

- Sample Preparation:

Prior to sampling, thoroughly mix stool specimen by vortexing for approximately 5 seconds or inverting (e.g., 3 - 4 times), or with a wooden applicator stick for formed/solid specimens. Remove the yellow top cap from the Aioprep device and set aside.

## Note: Prepared specimens in the Aioprep / Sample Diluent may be held at room temperature (19 - 27 C) for up to 8 hours or may be stored at refrigerated temperature (2 - 8 C) for up to 24 hours, prior to proceeding to step 2 for testing.

Note: Ensure that the measurement markings on the transfer pipette are not obscured during sample aspiration.

Pipettable, liquid/semi-solid specimens:

  - Human stool specimens, preserved in Cary-Blair or C&S media:

    - Using the provided transfer pipette, transfer 300 µL of preserved stool (third mark from pipette tip) directly into the Sample Diluent by inserting the tip of the pipette through the metering insert while expelling specimen to ensure the stool is transferred. NOTE: Transferring too little specimen, or failure to mix and thoroughly suspend the specimen in the Aioprep/ Sample Diluent, may result in a false-negative result. Refer to the examples below.  Recap the Aioprep device.

    - Proceed to step 2.

  - Human stool specimens, unpreserved:

    - Using the provided transfer pipette, transfer 100 µL of unpreserved stool (first mark from pipette tip) directly into the Sample Diluent by inserting the tip of the pipette through the metering insert while expelling specimen to ensure the stool is transferred. NOTE: Transferring too little specimen, or failure to mix and thoroughly suspend the specimen in the Aioprep/ Sample Diluent, may result in a false-negative result. Refer to the examples below. Recap the Aioprep device

    - Proceed to step 2.

Metering Device

User may encounter slight

resistance when inserting

tip of pipette through

metering insert.

Non-pipettable, formed/solid specimens (either preserved in Cary-Blair or C&S media or unpreserved):

- Using the Curian Campy sample collection brush, insert the brush into the stool specimen to coat the bristles of the brush. It is important that only the brush is covered with stool. Do not get stool on the shaft or handle of the brush; wipe any excess stool on the inside of the specimen container. Refer to the examples below.

## Correct	Incorrect- Under-sampling	Incorrect- Over-sampling

- Insert the Curian Campy sample collection brush into the Aioprep device, pushing the brush through the metering insert. Snap off the sample collection brush at the marked ‘Break’ point and discard the handle. Recap the Aioprep device, leaving the Curian Campy sample collection brush in the device.

- Proceed to step 2.

- Do not invert or shake Aioprep. Thoroughly mix the sample in the Aioprep sample diluent by vortexing for approximately 5 seconds. Sample must be thoroughly suspended in the Aioprep sample diluent. Remove the white tip cap from the bottom of the Aioprep and discard. While holding the Aioprep vertically, squeeze in the center of the barrel to add 3 drops into the ‘SAMPLE’ port of the test card. Discard the Aioprep device immediately.

- Start the Test Read by navigating to the ‘TEST’ Menu, select ‘ANALYZE NOW’ or ‘INCUBATE AND ANALYZE’, and then follow the on-screen instructions.

## For Analyze Now:

  - Incubate the test card at 19 - 27 C on the benchtop for 20 minutes.

  - Enter Sample ID and press Ok.

  - Within 2 minutes of the end of incubation, insert the test card into the drawer of the Curian Analyzer and close the drawer. Analysis of the reaction will be initiated by the Curian Analyzer.

  - The Curian Analyzer will analyze the test card and automatically report the test result.

## For Incubate and Analyze:

- Enter Sample ID and press Ok.

- Immediately insert the test card into the drawer of the Curian Analyzer and close the drawer. The analyzer will time the incubation for 20 minutes.

- After the end of the incubation, the Curian Analyzer will analyze the test card and automatically report the test result.

This facility’s procedure for specimen preparation is:____________________________  ______________________________________________________________________

This facility’s procedure for sample labeling is:_________________________________ ______________________________________________________________________

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Result interpretation is completed automatically by the Curian Analyzer system. The result will be shown on screen. Results can be retrieved from analyzer storage, printed and/or exported.

Patient Test Results:

One of the following result interpretations will be generated by the Curian Analyzer for patient specimens.

- Positive: Campylobacter antigen is present in sample.

- Negative: Campylobacter antigen is not present in sample.

- Control Invalid: Result indicates inadequate flow of sample. Repeat the test from the original stool specimen. If repeat testing yields the same results, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance.

External Quality Controls:

One of the following result interpretations will be generated by the Curian Analyzer for External Controls.

- Pass: This indicates the test card and reagents are performing as intended.

- Fail: This result indicates the test card and/or reagents are not performing as intended or a user error has occurred. Test should be repeated to assist in trouble-shooting the error. If repeat testing results in a failed output, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance. Do NOT perform patient testing if either of the QC test results are incorrect.

- Control Invalid: This result indicates inadequate sample flow. Repeat the test. If repeat testing yields the same results, please contact Meridian’s Technical Services Department at 1-800-343-3858 (US) or your local distributor for additional assistance.

Instrument Check:

One of the following result interpretations will be generated by the Curian Analyzer for Instrument Check.

- Pass: This indicates that the essential and critical components of the Curian Analyzer are working correctly.

- Fail: This indicates that there may be an issue with the Curian Analyzer. Refer to the Curian Operator’s Manual if an error message has occurred. Repeat testing to assist in trouble-shooting the error. If repeat testing results in a failed output, please contact Meridian’s Technical Services Department at 1-800-343- 3858 (US) or your local distributor for additional assistance.

In the event this test becomes inoperable, this facility’s course of action for patient samples is:_______________________________________________________________________

REPORTING OF RESULTS:

Negative test:  Report test results as Campylobacter antigen is not present in sample.

Positive test:  Report test result as Campylobacter antigen is present in sample.

EXPECTED VALUES:

The prevalence of Campylobacter infection observed during the 2020 study period for prospectively collected specimens was 1.2% (18/1474). The prevalence of Campylobacter spp. by sex assigned at birth for the tested population is provided below. The age demographic for the tested population ranged from less than 1 year old to 99 years old, with an average age of 48 years old.

| Sex Assigned at Birth | Samples Positive by Curian Campy and Reference Method | Prevalence in Tested Population |
| --- | --- | --- |
| Males (n=578) | 10 | 1.7% |
| Females (n=896) | 8 | 0.9% |
| Overall (n=1474) | 18 | 1.2% |

ASSAY SPECIFICITY:

CROSSREACTIVITY / MICROBIAL INTERFERENCE:

The Curian Campy assay was evaluated for cross-reactivity and microbial interference with the organisms listed below. Unless otherwise indicated, each organism was tested at minimum concentrations of 1.0 x 107 CFU/mL for bacteria/fungi or 1.0 x 105 TCID50/mL for viruses. The assay’s reactivity with Norovirus was evaluated using clinical Norovirus-positive stool specimens. None of the organisms showed cross-reactivity or microbial interference in the Curian Campy assay, except for Campylobacter helveticus, which was found to be positive at concentrations greater than 3.75 x 106 CFU/mL in unpreserved stool and 7.50 x 106 CFU/mL in preserved stool.

| Acinetobacter baumannii | Klebsiella pneumoniae |
| --- | --- |
| Aeromonas hydrophila | Lactobacillus acidophilus |
| Bacillus cereus | Lactococcus lactis |
| Bacillus subtilis | Listeria monocytogenes |
| Bacteroides fragilis | Peptostreptococcus anaerobius |
| Campylobacter concisus | Plesiomonas shigelloides |
| Campylobacter fetus | Porphyromonas asaccharolytica |
| Campylobacter helveticus | Prevotella melaninogenica |
| Campylobacter hyointestinalis | Proteus vulgaris |
| Candida albicans | Pseudomonas aeruginosa |
| Citrobacter freundii | Pseudomonas fluorescens |
| Clostridium bifermentans | Salmonella enterica susb. enterica serovar Hilversum |
| Clostridium difficile | Salmonella enterica susb. enterica serovar Typhimurium |
| Clostridium perfringens | Salmonella enterica susb. enterica serovar Minnesota |
| Edwardsiella tarda | Serratia marcescens |
| Enterobacter cloacae | Shigella boydii |
| Enterococcus faecalis | Shigella dysenteriae |
| Escherichia coli | Shigella flexneri |
| Escherichia coli EIEC | Shigella sonnei |
| Escherichia coli EPEC | Staphylococcus aureus |
| Escherichia coli ETEC | Staphylococcus aureus (Cowan’s) |
| Escherichia coli O157:H7 (non-toxigenic) | Staphylococcus epidermidis |
| Escherichia coli O157:H7 (toxigenic) | Streptococcus agalactiae |
| Escherichia fergusonii | Streptococcus dysgalactiae subsp. equisimilis |
| Escherichia hermanii | Vibrio parahaemolyticus |
| Helicobacter pylori | Yersinia enterocolitica |
| Adenovirus Type 1, 2, 3, 5, 40, 41 | Human Coronavirus |
| Coxsackievirus B2, B3, B4, B5 | Human Rotavirus |
| Echovirus 9, 11, 18 | Norovirus |
| Enterovirus 68, 69, 70, 71 | Parechovirus 1 (formerly Echovirus 22) |

TEST FOR INTERFERING SUBSTANCES:

The chemical and biological substances listed below were evaluated at the indicated concentrations for interference in the Curian Campy assay. None of the substances showed interference with the Curian Campy assay performance.

| Barium Sulfate (5% w/v) | Mylanta® (4.2 mg/mL) |
| --- | --- |
| Benzalkonium chloride (1% w/v) | Naproxen sodium (5% w/v) |
| Ciprofloxacin (0.25% w/v) | Nonoxynol-9 (1% w/v) |
| Ethanol (1% w/v) | Nystatin (1% w/v) |
| Hog gastric mucin (3.5% w/v) | Palmitic Acid/Fecal Fat (40% w/v) |
| Human blood (40% v/v) | Pepto-Bismol® (5% v/v) |
| Human hemoglobin (10.0% w/v) | Phenylephrine (1% w/v) |
| Human urine (5% v/v) | Polyethylene glycol 3350 (10% w/v) |
| Hydrocortisone (1% w/v) | Prilosec OTC® (5 µg/mL) |
| Imodium® A-D (5% v/v) | Sennosides (1% w/v) |
| Kaopectate® (5% v/v) | Simethicone (10% w/v) |
| Leukocytes (0.05% v/v) | Stearic Acid/Fecal Fat (40% w/v) |
| Mesalazine (10% w/v) | Tagamet® (5 µg/mL) |
| Metronidazole (0.25% w/v) | TUMS® (50 µg/mL) |
| Mineral Oil (10% w/v) | Vancomycin (0.25% w/v) |

LIMITATIONS OF THE PROCEDURE:

- The Curian Campy assay must only be used with the Curian Analyzer.

- Curian Campy is a qualitative, in vitro diagnostic test. The Curian Analyzer will only provide qualitative results. This test is not intended to provide quantitative results.

- Curian Campy detects both viable and nonviable Campylobacter bacteria and may yield a positive result in the absence of living organisms.

- A negative test result does not definitively rule-out the presence of Campylobacter species in suspected patients. Levels of organism may be present in feces beneath the limit of detection for the Curian Campy assay, and therefore, if Campylobacter is suspected, alternative testing should be conducted.

- A negative result may occur if the sample was collected, transported, or stored improperly.

- Failure to follow the test procedure may adversely affect test performance and/or invalidate the test result.

- Test results must be evaluated in conjunction with other clinical data available to the physician.

- Negative test results do not rule out other possible infections.

- Positive test results do not rule out co-infections with other pathogens.

- Curian Campy is designed to be used with unpreserved stool samples or stools preserved in Cary-Blair or C&S media. DO NOT USE stool on swabs or in other transport media (e.g., formalin, polyvinyl alcohol). No data exists on the effects of colonic washes, barium enemas, laxatives, or bowel preparations on the performance of Curian Campy. All of these procedures can result in extensive dilution or the presence of additives that may affect test performance.

- Transferring too little specimen, or failure to mix and completely suspend the specimen in the Sample Diluent, may result in a false negative test result.

- The performance of this test has not been evaluated for use in patients without signs and symptoms of gastroenteritis.

- Campylobacter helveticus (strain ATCC 51209) at levels greater than 3.75 x 106 CFU/mL in unpreserved stool or greater than 7.50 x 106 CFU/mL in preserved stool may cross react or interfere with the performance of the test.

- Non-pipettable, formed/solid specimens should only be processed with the Curian Campy Sample Collection Brush (Meridian Bioscience, Inc. Catalog 11516).

## PERFORMANCE CHARACTERISTICS:

Refer to Directional Insert – Meridian Bioscience curian CAMPY

## REFERENCES:

Refer to Directional Insert – Meridian Bioscience curian CAMPY