## Sheet1
TEST SYSTEM: Curian Shiga Toxin - EHEC
Steps | Failure Mode | Causes | Means of Failure Detection
|  |  | Internal Controls | External Controls | Engineering Controls | Operator Training | Laboratory's Mitigating Actions
List the stage or aspect of the test system’s process under investigation. | List all manners in which failure could occur in this step. | List all causes of the failure mode that have the potential to produce incorrect test results. | Are there internal biological or procedural controls to detect failure? | Can external controls increase the probability to detect failure? | Are there manufacturer checks to reduce the probability of failure? | Can operators reduce or detect failures through training? | What other processes can the laboratory implement to reduce failure?
Specimen collection | Specimen degradation or specimen contamination | Improper Specimen type | No | No | No | Yes | Techs are trained that only stool or stool in Cary Blair type transport media are approved specimen types. Techs will reject specimens that are not considered stool including colonoscopy specimens, aspirates etc.
|  | Use of wrong type of transport media | No | No | No | Yes | Transport media vial is provided to the patinet along with collection, storage and transport instructions Nursing staff collects from inpatients and there is an SOP on proper collection and handling for all assays.
|  | Improper storage and transport | No | No | No | Yes | Transport media vial is provided to the patinet along with collection, storage and transport instructions Nursing staff collects from inpatients and there is an SOP on proper collection and handling for all assays.
|  | Interfering substances | Yes | No | No | No | No interfering substances were identified in manufacturer's studies. Any interference should also cause the internal control to fail.
Specimens and Reagents | Sample Processing | Insufficient mixing of stool specimen | No | No | No | Yes | Techs are trained to the proper amount of stool to add to the broth.
|  | Incorrect amount of stool added to the GN or MacConkey broth, then added to theAioprep device | No | No | No | Yes | 1. Incorrect amount of specimen may not cause an issue due to growth requirement in the broth.
2. Training and competency assessment must be completed and signed off prior to performing any testing where patient results are reported. 3. Competency is completed initially, than at 6 months and annually thereafter. 3. Eash user participates in proficiency testing.
|  | No visual growth in GN or MacConkey broth | No | No | No | Yes | Users are trained to check for visible groth in the broth and if no growth, the inoculation and incubation are repeated.
|  | Specimen or kit not warmed to room temperature | Yes | No | No | Yes | If specimen or kit are not warmed appropriately, the internal control may not develop.
Users are trained to get kit our of cooler 1 hour prior to testing. Users are trained initially, then after 6 months, and then annually.
|  | Incorrect amount of prepared specimen added to test card | Yes | No | No | Yes | Internal control monitors flow and insufficient specimen adding that impacts flow will be reported as invalid.
|  | Incorrect test card used | No | No | Yes | Yes | Analyzer reads the barcode and determines which assay is being run. Training and competency are completed and documented prior to results being reported.
|  | Incorrect analyze option chosen on instrument | Yes | No | No | Yes | Incorrect option chosen will not impact results.
Operator | Operator Capacity | Untrained operator | Yes | No | Yes | No | Internal control monitors processing of specimens. Only trained techs will have a user name and password to access the analyzer.
| Laboratory Staffing | Short staffing | No | No | No | No | In the event of short staffing, TAT may be impacted. Only trained operators can run the test and know the log in credentials.
|  | Operator staffing credentials | Yes | No | No | No | Hiring manager and human resources will verify all operators are qualified per local, federal and state guidelines for moderately complex testing.
Reagents | Test Kit Degradation | Shipping | Yes | Yes- All lots are tested prior to putting into use | No | Yes | Assay has an internal control that helps to confirm reagent integrity. External controls are tested on all new lots or shipments. Product is shipped in labelled boxes that detail required storage conditions. Operators will notify Laboratory Manager if there are any defects are noticed upon receipt of product.
|  | Storage | Yes | No | No | Yes | All refrigerators are rooms have temperature monitored and recorded daily. Trained operators store product according to the label on the kit box.
|  | Used past expiration | No | No | Yes | Yes | Expiration date is checked by the analyzer for each test card
| Curian analyzer failures | Optics | Yes | No | Yes | No | Self test automatically performed every 24 hours and upon turning on analyzer to detect issues. Instrument check card tested minimally every 30 days.
|  | Test card drawer not fully closed | Yes | No | Yes | Yes | Staff is trained to proper use of analyzer. A sensor detects when the drawer is closed.
|  | Software failures | No | No | Yes | No | Self test which is automattiically performed upon turing on the analyzer or every 24 hours will detect software issues. Lab will call Technical Support if a faiure occurs.
Laboratory Environment | Atmospheric Environment | Dust | No | No | No | Yes | Instrument is closed and only opened upon running a test. Benches are cleaned and it is documented daily.
|  | Temperature | Yes | No | No | Yes | Temperature of the lab is monitored and documented daily. Instrument operating temperature is 15 C to 35 C (59 F to 95 F).
Laboratory Environment |  | Humidity | No | No | No | Yes | Relative humidity 10% to 70% non-condensing. Humidity is monitored and documented daily.
| Utility Environment | Electrical | No | No | Yes | No | Analyzer requires installation of ferrites on accessories to prevent electrical interference. Power on self test verifies proper electrical requirements are met. Failed test locks operators from using instrument.
| Transcription Errors | Incorrect results input for patient | No | No | No | Yes | Results are spot checked for accuracy prior to reporting. Log is kept of all results that are corrected. Once a month, an audit is performed on results reported.
| Timely reporting of results | results delayed | No | No | No | Yes | Pending list printed twice per shift. Techs trained to report results upon completion of the test. All results are reported before leaving the lab for the day. No results are left to be reported on the next shift.